Attachment 4 – Equipment & Reagent Requirements

L.3.1.1 - DATA MANAGEMENT SYSTEM (Section C – Statement of Work, C.6.2 Data Management System

The offeror shall identify the data management system that meets the requirements.

| Name of Data Management System | Point of Contact |
| --- | --- |

| INSTRUCTIONS: The contractor shall provide a data management system for drug testing with a user interface that accepts specimen identification and test specifications. The data management system shall be capable of storing the data elements specified in Attachment 3 - Specifications for Probation and Pretrial Case Tracking System (PACTS) Drug Testing Results (DTR) component On-Site Laboratory(s) Interface and in accordance with SOW Section C.6. Indicate by YES or NO if you are capable to meet the requirements below: |  |  |  |
| --- | --- | --- | --- |
| ITEM # | Performance Requirement | Meets Requirement |  |
|  |  | Yes | No |
| 1 | The contractor shall provide a web-based data management system containing, at a minimum, all database fields, elements, functions, and reports described in Attachment 3 and Section C.6.4. |  |  |
| 2 | Fully compatible with, and capable of interfacing with all analyzers and equipment proposed and used in each on-site laboratory. |  |  |
| 3 | The contractor shall provide all required hardware—including CPU, monitor, printer, bar-code scanner, bar-code labels, and bar-code printer—and all necessary software for each workstation within the on-site laboratory. |  |  |
| 4 | Provide comprehensive training to all on-site laboratory staff on the use of the data management system, including specimen entry, report generation, and troubleshooting. |  |  |
| 5 | The contractor shall identify a dedicated point of contact to assist on-site laboratories with any data-management-system support, questions or problems. |  |  |
| 6 | The contractor shall provide administrative access rights to specified users. |  |  |
| 7 | On-Site Modifications: Enter new vendor locations, add/delete Officers or collectors, add/delete tests when creating a requisition, Correct donor info, delete duplicate requisitions, and merge duplicate clients. |  |  |
| 8 | The data management system shall: - Electronically record the results of each assay. - Read bar-coded specimens. - Include fields for specimen ID number and special instructions. - System shall transmit individual test results—quantitative when appropriate—to remote locations within two (2) hours of acceptance of the test results. - Import calibration and quality-control values from the analyzer. - Produce customized aggregate reports in user defined categories |  |  |

L.3.1.2 DRUG LABORATORY EQUIPMENT (Section C – Statement of Work, C.5.2.1 – Performance of Equipment)

| INSTRUCTIONS: The Offeror shall provide all the required equipment and supplies, including but not limited to calibrators, controls, and consumables, to meet the testing needs of each on-site laboratory. The contractor shall provide equipment that is capable of testing specimens based on the following volume and testing requirements. Indicate by YES or NO if you are capable to meet the requirements below: |  |  |  |  |
| --- | --- | --- | --- | --- |
| ITEM # | Performance Requirement | Meets Requirement |  |  |
|  |  | Yes | No | Name of instrument and testing capacity. Page # in proposal. |
| 1.a | No less than 100 tests (test represents an individual analyte, e.g. cocaine) per hour for laboratories conducting between 500 and 4,999 tests per month; |  |  |  |
| 1.b | Supported by Manufacturers equipment |  |  |  |
| 2.a | No less than 400 tests (test represents an individual analyte) per hour for laboratories conducting between 5,000 and 11,999 tests per month; |  |  |  |
| 2.b | Supported by Manufacturers equipment |  |  |  |
| 3.a | No less than 800 tests (test represents an individual analyte) per hour for sites conducting between 12,000 and 20,000 tests per month. |  |  |  |
| 3.b | Supported by Manufacturers equipment |  |  |  |
| 4 | Laboratories conducting over 20,000 tests (test represents an individual analyte) per month, the offeror shall provide an additional analyzer capable of testing no less than 800 tests per hour. |  |  |  |

L.3.1.3 REAGENTS (Section C – Statement of Work, C.5.5.1 Reagent Chart)

| INSTRUCTIONS: The Offeror shall confirm they can provide ready-to-use liquid reagents to use to test for specified drugs of abuse at the specified cutoff levels set below. Indicate by YES or NO if you are capable to meet the requirements set forth below: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| REAGENT DESCRIPTION | CUTOFF LEVELS | Meets Requirement for Cut-Off Level |  | Meets Requirement for Liquid Ready to Use |  |  |
|  |  | Yes | No | Yes | No | Manufacturer of Reagent. Page # in proposal |
| 1) Cannabinoids | 50 ng/mL |  |  |  |  |  |
| 2) Cocaine Metabolite | 150 ng/mL |  |  |  |  |  |
| 3) Opiates (Based on 300 ng/mL of morphine or other equivalent opiate) | 300 ng/mL |  |  |  |  |  |
| 4) Amphetamines (Based on 500 ng/mL d-methamphetamine and/or d-amphetamine, but should be able to detect amphetamine, MDMA and MDA at clinical concentrations) | 500 ng/mL |  |  |  |  |  |
| 5) Benzodiazepines (Based on 200 ng/mL of oxazepam or another equivalent benzodiazepine) | 200 ng/mL |  |  |  |  |  |
| 6) Creatinine | Normal: >20 mg/dL |  |  |  |  |  |
|  | Dilute: 2-20 mg/dL |  |  |  |  |  |
|  | Substituted: < 2mg/dL |  |  |  |  |  |
| 7) Methadone | 300 ng/mL |  |  |  |  |  |
| 8) Oxycodone | 100 ng/mL |  |  |  |  |  |
| 9) Buprenorphine | 5 ng/mL |  |  |  |  |  |
| 10) Barbiturates | 200 ng/mL |  |  |  |  |  |
| 11) 6 AM | 10 ng/mL |  |  |  |  |  |
| 12) Fentanyl | 1 ng/mL |  |  |  |  |  |
| 13) Ethyl Glucuronide (EtG) | 500 mg/mL |  |  |  |  |  |
| 14) Ketamine | between 50 - 100 ng/mL |  |  |  |  |  |

L.3.1.4 KEY PERSONNEL (Section C – Statement of Work, H.3 Key Personnel)

| Name of Key Personnel | Position | Meets Requirements | Resume Included |
| --- | --- | --- | --- |