Salient Characteristics:

Image and Data Management System must be FDA 510k approved and listed as a Medical Device

Image and Data Management System must provide “Post Processing Analysis,

Dynamic Review/Manipulation of RAW data” functionality for ZEISS Cirrus

HD-OCT. If Glaucoma Workplace is procured you can do with the ZEISS

HFA Instruments as well.

Image and Data Management System must be able to store and access raw

instrument data from ZEISS Field Analyzer, ZEISS Cirrus HD-OCT, ZEISS

Visucam Fundus Camera and ZEISS FF450/Visupac Fundus Camera systems.

Image and Data Management System must have functionality to enable

real-time and dynamic editing/manipulation of all prior Visual Field Test data

for all Glaucoma patients and allow for Baseline changes, real-time report

generation for any VFI date, display RealEye images and data from the new

ZEISS HFA 3 instruments.

Image and Data Management System must support DICOM standards and

Work-List functionality in order to create and send patient demographics to a

specific or all instrument’s via a DICOM work-list within the eye clinic’s

network.

Combined Report: HFA-Cirrus reports for Structure Function analyses in

Single Display with OD-OS on one (1) page print

out from data of Cirrus HD OCT and a Humphrey Visucal Field. If Combined

Report is procured as option of FORUM.

Image and Data Management System must have Bi-directional DICOM

capabilities to send both DICOM Modality Work-lists containing

Consult/Order details to instrument’s within the eye-clinics network and

receive back the diagnostic test results and RAW data for storage within the Image and Data Management System for access.

Image and Data Management System must connect to any brand of

non-DICOM diagnostic instrument with export capabilities, and convert the

diagnostic test results into a DICOM compatible file containing all patient

demographical information, UID’s, Accession number, Device name and type,

laterality and other key pieces of data with the least amount of user interaction.

This must be then stored within the Image and Data Management System and

also forwarded to the NAVY ECIA (Enterprise Clinical Imaging Archive)

Image and Data Management System must not require the use of an external

PC or import computer for a DICOM compatible modality.

Image and Data Management System must be able to download all past RAW

data from one or more ZEISS Field Analyzers, ZEISS Cirrus HD-OCT’s,

ZEISS Visucam Fundus Camera and ZEISS FF450/Visupac Fundus Cameras

acquired within the main clinic and all satellite clinics upon installation of the

Image and Data Management System.

Image and Data Management System must be able to display ZEISS Field

Analyzer Glaucoma Progression Analysis (GPA) and enable the clinician to

manipulate data and store baseline changes at the workstation within the

Image and Data Management System if Glaucoma Workplace is procured as a

part of FORUM

Image and Data Management System must be able to display and allow the

clinician to scroll through the various ZEISS Cirrus Macula Cube B-scan

slices, within the Image and Data Management System viewer.

Image and Data Management System must be able to display images in

pre-formatted clinical display groupings for Glaucoma, Retina and Cataract

cases.