Revised Draft: Call for Solutions

Call for Solutions (CFS) Number: W911QY-26-S-PPTS-01
Issuing Office: U.S. Army Contracting Command-Aberdeen Proving Ground (ACC-APG), Natick Contracting Division (NCD)

Date Issued: 12 March 2026

TITLE: Portable Patient Transport System (PPTS)

SECTION A: GENERAL INFORMATION

- This Call for Solutions (CFS) is issued under the authority of 10 U.S.C. § 4022, "Authority of the Department of Defense to carry out certain prototype projects", against the Army Open Solicitation (AOS) W9128Z-25-S-A002 under BAA authority.

- The Government intends to award multiple Prototype Other Transaction (OT) Agreements to the offerors whose Solution Briefs are determined to be most advantageous to the Government based on the evaluation criteria outlined in Section D.

- This CFS invites interested parties to submit a "Solution Brief" for a prototype PPTS. Based on the evaluation of Solution Briefs, the Government may issue a Request for Solution Proposal (RSP) to selected offerors for a more detailed submission.

SECTION B: BACKGROUND / PROBLEM STATEMENT / CAPABILITY GAP

The Program Executive Office for Chemical, Biological, Radiological, and Nuclear Defense (PEO-CBRND) and Joint Project Manager for CBRN Protection are seeking solutions for a single, non-ambulatory patient system to support the Biocontainment, Isolation, and Quarantine (BIQ) PPTS program. This man-portable system will control and prevent the spread of biological threats (biothreats) while transporting a single, non-ambulatory infectious patient by air, rail, ground, or sea.

The Department of War has an urgent need for a Portable Patient Transport System (PPTS) to provide safe, contained, and regulated intra- and inter-theater transport of patients with highly infectious diseases. The system must be certified for flight on multiple military airframes, be capable of being transported in select ambulances, and be operable and sustainable in austere environments. The system must protect the force, the platform, and the crew from contamination while allowing a high standard of medical care to the patient.

This effort will follow a multi-phased acquisition strategy as detailed in the Statement of Objectives (SOO) and Section E of this CFS. Key objectives include achieving airworthiness certification, ensuring robust bio-containment, enabling long-duration patient sustainment, and acquiring vendor FDA 510(k) clearance. A detailed breakdown of all performance requirements is provided in the associated SOO, AOS Statement of Objectives PPTS_.docx.

SECTION C: SUBMISSION INSTRUCTIONS

- Initial Submission: Solution Brief

- Submission Format: The total submission shall consist of two (2) separate files:

  - File 1: Solution Brief

  - File 2: Supporting Data Package

- File 1: Solution Brief

  - The Solution Brief is the core of the offeror's proposal. In accordance with the parent Army Open Solicitation (AOS), the Solution Brief is limited to fifteen (15) pages.

  - This page limit includes the cover page, table of contents, and the completed CFS Attachment 0001, "PPTS Capabilities Assertion Matrix."

  - The Solution Brief should provide a concise narrative addressing the technical approach, innovation, IP strategy, and business information. The Assertion Matrix must be fully completed, with the "Supporting Data Reference" column clearly pointing to the specific location of the evidence in File 2 This submission will only count as a one (1) page portion. of submission within the solution brief submission.

- File 2: Supporting Data Package

  - The Supporting Data Package contains all the evidence and verifiable data referenced in the Assertion Matrix.

  - There is no page limit for the Supporting Data Package.

  - This package should be clearly organized (e.g., with a table of contents and bookmarks) to allow for easy navigation and verification of claims.

  - Acceptable supporting data includes but is not limited to: third-party test reports from accredited laboratories, detailed engineering analyses, modeling and simulation data, white papers with supporting calculations, videos of internal testing, or existing commercial/government certifications. The burden of proof rests with the offeror.

- Submission Deadline & Method:

  - Both files shall be submitted in PDF format no later than [DATE TBD] at 1600 EST.

  - Submissions shall be sent via email to [TBD email address]. The email subject line shall be: "CFS W911QY-26-S-PPTS-01 Submission - [Offeror Name]".

- The two files shall be named as follows:

  - [Offeror Name] PPTS_Solution_Brief.pdf

  - [Offeror Name] PPTS_Supporting_Data.pdf

The initial solution briefs MUST assert to meeting or exceeding the stated Threshold Requirement to be considered for a request for proposal submission for Phase 1. Support of assertions may be included in Supporting Data file, to be excluded from the page limitation within the Solution brief.

|  | Threshold Requirement | Desired Requirements | Validation Method | Evaluation Criteria Against Threshold |
| --- | --- | --- | --- | --- |
| Force Protection | Negative Pressure Differential. 0.01 inches of water gauge. | The PPTS shall maintain a negative pressure differential of 0.01 inches of water gauge or more inside the containment enclosure with all pass-through ports sealed and the transfer pouches closed, if applicable. | Government Analysis of assertion support/ Demonstration | This is a pass/fail criterion |
|  |  | The PPTS shall maintain a negative pressure differential of 0.01 inches of water gauge or more inside the containment enclosure for 2 minutes with one large pass-through port open and the transfer pouches closed, if applicable. |  |  |
|  | Air Flow Direction. Head to Foot. | Same as Threshold. | Government Analysis of assertion support/ Demonstration | This is a pass/fail criterion |
|  | Air Exchanges. 12 Exchanges / Hour. | Same as Threshold. | Government Analysis of assertion support/ Demonstration | This is a pass/fail criterion |
|  | Filtration Efficiency. 99.97% of 0.3 micron particles. | The PPTS shall have an intake air filter with a minimum filtration efficiency of least 99.97% at 0.3-micron particle diameter. | Government Analysis of assertion support/ Demonstration | This is a pass/fail criterion |
|  |  | The PPTS shall have an exhaust air filter with a minimum filtration efficiency of 99.97% at 0.3-micron particle diameter. |  |  |
|  |  | The PPTS intake filters shall have a removable device to prevent ingestion of rain. |  |  |
|  |  | The PPTS intake filter shall be located outside of the containment enclosure. |  |  |
|  |  | The PPTS exhaust filter shall be located outside of the containment enclosure. |  |  |
|  |  | The PPTS exhaust filter shall be upstream of the blower. |  |  |
|  |  | The PPTS air filtration system shall be able to operate in rain conditions without performance degradation. |  |  |
|  | Filter Certification. National Institute for Occupational Safety and Health (NIOSH) P-100 certified filters. | Same as Threshold. | Government Analysis of assertion support/ Demonstration | This is a pass/fail criterion |

Additionally, the Government will assess whether the offeror meets one of the conditions under 10 U.S.C. § 4022(d)(1), with a focus on significant participation from non-traditional defense contractors and/or small businesses. Offerors shall provide support to their approach to meeting one of the conditions at 10 U.S.C. § 4022(d)(1).

Call for solution briefs should include:

Contractors shall include test plans and reports and any other relevant documentation to support claimed performance characteristics including maintaining negative pressure, filtration, patient and equipment access, and operational size and weight. In addition, contractors shall disclose any unique test methods required to assess the performance of their solution.

The contractor will use the PPTS SOO as a basis for preparing a draft Statement of Work (SOW), which will be included as an integral part of the proposal.  The contractor shall propose a delivery schedule in the SOW to include the number of systems that can be produced per month. The contractor shall also provide logistic details including packaging, handling, storage, and transportation requirements, as well as logistical footprints and shelf-life estimates

The Government requires sufficient data rights to operate, maintain, and sustain the PPTS. The Government’s strong preference is for Unlimited Rights for all technical data, software, and associated intellectual property developed under this effort. Special consideration will be given to IP strategies that facilitate the rapid and affordable achievement of military airworthiness certification. As part of this effort, the Government requires delivery of a complete Bill of Materials (BOM) for all prototype systems. The Government will evaluate the impact of any proposed data rights restrictions, and proposals that hinder future competition for sustainment may be evaluated negatively.

The contractor shall include in a table form, its data rights assertions for all technical data and computer software (including non-commercial technical data and non-commercial computer software as well as commercial technical data and commercial computer software) being provided to the Government with less than an unlimited or unrestricted rights license and explain how the approach will result in the best approach, considering life cycle costs.  The Government will consider the impact of the data rights assertions in making selection decision(s).  Please note that the Government reserves the right to evaluate any proposed approach that provides less than Government Purpose Rights (GPR) for the prototype design documentation and/or does not facilitate competitive procurement of hardware maintenance/sustainment services for PPTS, described herein, negatively and use that information as part of its selection decision(s).

The contractor shall provide the cost estimate in a format that allows for clear evaluation and comparison, ensuring completeness and accuracy. If possible, include assumptions, methodologies, and any relevant supporting documentation used to derive the estimate.

- Production Costs

  - Estimated costs for full-scale production, including materials, labor, and compliance with FDA 510(k) clearance and Berry Amendment requirements.

- Life-Cycle Cost Estimate (LCCE)

  - A comprehensive assessment of the system's affordability and long-term sustainment costs, including maintenance, operational support, and disposal.

- Additional Costs

  - Any other costs associated with meeting the requirements outlined in the Statement of Objectives (SOO) and achieving successful completion criteria for follow-on production.

Technical maturity, item cost and performance will be assessed during submission evaluation.

Follow-on Submission: Request for Solution Proposal (RSP)

  - Offerors whose Solution Briefs are selected will be invited to submit a full proposal in response to an RSP.

  - The full proposal will require a comprehensive submission, including but not limited to:

- A draft Statement of Work (SOW).

- A detailed Master Delivery Schedule for all hardware and data deliverables.

- The total firm-fixed-price for all prototypes and associated deliverables.

- The Performer’s projected unit cost and maximum monthly production rate for a potential Full Rate Production (FRP) quantity.

- A complete Bill of Materials (BOM) down to the raw material level.

- A table detailing the Performer’s data rights assertions for any technical data or software with less than Government Purpose Rights. The Government’s preference is for GPR.

- Supporting documentation for all performance claims, including but not limited to: test plans and reports, biological protection data, patient care (e.g., glove ports, visibility, waste storage, patient restraint) maintainability/reliability data, and platform integration analysis.

- Disclosure of any unique or non-standard test methods required to assess the proposed solution.

- Details on any special logistical requirements (e.g., decontaminability, maintainability, reliability, size, weight, power, operating environmental conditions, setup time, platform integration, and storage) and estimated shelf-life/service-life.

SECTION D: EVALUATION PROCESS & CRITERIA

The Government will evaluate submissions using a phased approach. The initial Request for Solution Proposals (RSP) will be assessed to select the most promising solutions for a Segment I award.

Part 1: Initial Downselect (RSP to Segment I Award)

Offerors' RSP will be assessed to select up to five (5) of the most promising solutions for a Segment I OT Agreement award. To be considered, the RSP must demonstrate a credible approach to meeting the project's core objectives. In addition to the affordability of the prototype and availability of fund, the following criteria will be used to determine the solutions to be most advantageous to the Government :

| Evaluation Area | How Assessed |
| --- | --- |
| 1. Technical Merit | The Government will evaluate the RSPs assertions with regards to meeting the Threshold Requirements. The Government will further evaluate the credibility and feasibility of the proposed technical approach to meeting the full PPTS requirements, with a specific focus on the proposed path to airworthiness certification and vendor-acquired FDA 510(k) clearance. |
| 2. Innovation & IP | The Government will assess the novelty of the proposed solution and the offeror's proposed Intellectual Property approach, favoring flexible strategies that support the Government's long-term needs for sustainment and certification. |

3. Affordability	The Government will assess the reasonableness of the estimated costs for full-scale production will be evaluated. This includes the completeness and accuracy of the estimate, as well as the inclusion of all assumptions, methodologies, and relevant supporting documentation used to derive the estimate.

Part 2: Downselect from Segment I to Segment II

Evaluation criteria for potential selection for the Segment II option will be based on the affordability of the prototype, availability of funds and the prototypes ability to meet the threshold requirements and user feedback in Segment I testing

At the conclusion of Segment I, the Government will select up to two (2) performers to proceed to Segment II based on the affordability of the prototype, availability of funds, asessment of their progress, prototype performance against the requirements in the SOO (Attachment 1, Tables 2 & 3), and demonstrated potential.

| Evaluation Area | Point of Measure |
| --- | --- |
| 1. Technical Feasibility & Requirements Alignment | Results from government-witnessed early system demonstrations and Segment I requirements testing; submitted Technical Data Package (TDP); technical reports. |
| 2. Integration Potential & FDA 510(k) Clearance Progress | Draft Interface Control Document (ICD); submission of the FDA 510(k) clearance request and performer's status with the FDA 510(k) clearance process. |
| 3. Design Maturity & Risk Reduction | Submitted Risk Management Plan and associated risk register; results from component/subsystem qualification testing; material selection reports. |
| 4. Programmatic Performance | On-time delivery of all Contract Data Requirements List (CDRL) items; adherence to the Integrated Master Schedule (IMS); performance in formal reviews. |

** After receipt of the lifecycle cost estimate, the Government will perform a detailed analysis of the submitted lifecycle estimated proposal. For purposes of this analysis, the Agreement Officer and/or a representative of the Agreement Officer shall have the right to examine the supporting records and/or request additional information as needed.

Part 3: Final Candidate Selection (End of Segment II)

At the end of all Segment II activities (which includes the completion of both Phase II and Phase III), the Government will perform a final, comprehensive evaluation. This evaluation will lead to a formal, unilateral decision on whether the prototype meets the Phase III Objectives, whether the prototype has been "successfully completed" as required by 10 U.S.C. § 4022(f), and will determine which candidate has performed best when assessed against the key technical goals in the evaluation areas below. This evaluation will determine the final candidate for any non-competitive follow-on production award.

This "successful completion" determination is a mandatory prerequisite for awarding any non-competitive follow-on production contract.

Phase III Objectives

To proceed to the final evaluation, the prototype must first meet the following objectives during Phase III:

- Achieve a minimum Technology Readiness Level (TRL) of 8 and a Manufacturing Readiness Level (MRL) of 8.

- Support Government-led flight testing on operational aircraft to validate performance in an aeromedical evacuation mission profile.

- Provide all necessary data and support to achieve a formal Airworthiness Certification.

- Acquire FDA 510(k) clearance.

The Government's assessment of "best performance" will be based on the following evaluation areas, which directly incorporate the objectives from Phase III:

| Evaluation Area | Point of Measure |
| --- | --- |
| 1. Validated Requirement Performance | The final validated Government Test Report from all testing phases, assessed against the requirements in the SOO. |
| 2. Regulatory Compliance and Airworthiness Data Maturity | Final approval of FDA 510(k) clearance; delivery of a complete and acceptable Airworthiness Data Package per MIL-HDBK-516C. |
| 3. Manufacturability & Production Readiness | Evidence of Manufacturing Readiness Level (MRL), the final Bill of Materials (BOM), and an overall production risk assessment. |
| 4. Affordability & Life-Cycle Cost | The performer's final production proposal and Life-Cycle Cost Estimate (LCCE) compared against the government's independent estimate. Less is best. |

SECTION E: ANTICIPATED PHASED EFFORT

Evaluation Criteria. At each segment of testing the government plans to make downselect decisions as to which proposals will move on to the next segment. In each phase the agency will evaluate the proposals in their ability to meet key technical goals and criteria in accordance with the table below. Downselect decisions will be made based on meeting the primary criteria first and then secondary criteria, if applicable. The proposals who meet all the primary criteria for each phase and are assessed as the best in the secondary criteria, where applicable, will be selected to move on to the next phase. A trade off analysis between evaluation criteria will be considered.

The Government intends to enter negotiations for the follow-on production effort with all remaining participants following Phase II down selection with the full understanding that any potential award can only be made after successful completion and is expected to occur after the completion of Phase III Testing and final down selection.

- Segment I (Phase I): Initial System Delivery & Evaluation

  - Award of up to five (5) OT agreements.

  - Deliverable: Thirty (30) prototype systems per awardee.

  - Goal: Achieve minimum Technology Readiness Level (TRL) 5 / Manufacturing Readiness Level (MRL) 5. Performers will develop initial designs and support Government-led testing.

  - Downselect to two (2) systems at the end of Segment I (Phase I)

- Segment II (Phases II & III): Advanced Integration & Certification

  - Deliverable: Ten (10) additional prototype systems per awardee.

  - Phase II Goal (TRL 6 / MRL 6): Focus on hardware integration with aircraft and ambulances, environmental testing, and RAM data collection.

  - Phase III Goal (TRL 8 / MRL 8): Focus on operational flight testing and maturation of the data package to achieve formal Airworthiness Certification and FDA 510(k) clearance.

- Follow-on Production

  - Upon successful completion of the prototype project by the final selected performer or performers, the Government may award a non-competitive Follow-on Production OT or FAR-based contract, in accordance with 10 U.S.C. § 4022(f), for up to 400 units over four years.

SECTION F: INTELLECTUAL PROPERTY

The Government’s strong preference is for Unlimited Rights for all technical data, software, and associated intellectual property developed under this effort. Special consideration will be given to IP strategies that facilitate the rapid and affordable achievement of military airworthiness certification. As part of this effort, the Government requires delivery of a complete Bill of Materials (BOM) for all prototype systems. The Government will evaluate the impact of any proposed data rights restrictions, and proposals that hinder future competition for sustainment may be evaluated negatively.

SECTION G: SUCCESSFUL COMPLETION CRITERIA FOR FOLLOW-ON PRODUCTION

The Government will make a formal determination of "successful completion" for the final selected prototype.

Criteria for "Successful Completion"

A prototype project is considered "successfully completed" only after a written determination is made by the responsible approving official. This determination must state that the prototype has either:

- Met the stated Threshold Requirements; and satisfied success metrics incorporated into the Prototype OT

OR

- Accomplished a particularly favorable or unexpected result that justifies a transition to production.

To be deemed "successfully completed" and therefore be eligible for a non-competitive follow-on production award, the prototype must be judged by the Government to have met all minimum threshold requirements outlined in the SOO. The key criteria that will inform the Government's final determination include, but are not limited to, the following, as verified by government-conducted or government-witnessed testing and analysis:

| Requirement | Threshold |
| --- | --- |
| Leak Containment | The PPTS shall hold at least 3 liters of liquid contamination inside the containment enclosure without leaking. |
| Force Protection | Negative pressure differential of 0.01 In.w.g |
|  | Head to foot airflow |
|  | 12 Air Changes per Hour (ACH) |
|  | 99.97% at 0.3-micron particle diameter |
| Power | 12 hours continuous operation |
|  | System shall provide a low battery warning |
|  | During battery hot swap, power interruption shall be less than or equal to 1 minute |
| Regulatory Compliance | System shall have an FDA 510(k) clearance |
| Occupant/Crew Interface | System must provide visibility of the occupant from head-to-waist |
|  | System will have pass-through interfaces from the system exterior to the system interior |
|  | System will include glove ports located near patient head, torso, and feet to enable proper patient care |
| Air Filtration System | The system must be able to indicate system interruption via alarms |
| Weight | System will comply with the person-portable lifting limits specified in MIL-STD-1472 |
| Interoperability | System shall accommodate a 90% of adult male and females (5th to 95th percentile) |
| Airworthiness | Collection of test data including favorable results for EMS, altitude, rapid decompression, crash hazard acceleration, explosive atmosphere, and NVD testing. |
|  | The System must meet all military, federal, and international airworthiness and transport requirements for airlift in accordance with: MIL-HDBK-516 for C-17, C-130, KC-135 and HH-60. |
| Patient Restraint | The PPTS restraint straps shall be capable of being released without breaking containment regardless of the state of attachment to a conveyance and/or transport platform. |
|  | The PPTS shall have restraint straps for safe physical patient restraint inside the containment enclosure appropriate to the manner of movement regardless of state of attachment to a vehicle. |
|  | The PPTS restraint straps shall accommodate 90% of personnel for all key sizing dimensions (5th to 95th percentile, male and female body dimensions). |
|  | The PPTS restraint straps shall be in the vicinity of the patient’s chest, waist, and legs. |
| Transfer Pouch | The PPTS shall provide at least one transfer pouch to allow items to be passed into the containment area while maintaining a negative pressure differential of 0.01 inches of water gauge or more inside the containment enclosure during the transfer. |
| Exterior Disinfection | The PPTS exterior surfaces shall be compatible with one or more Environmental Protection Agency (EPA)-registered disinfectants for a contact time of 15 minutes. |
| Litter Compatibility | System shall be compatible with NATO litter |
| Ambulance Integration | System shall be compatible with ground vehicles UFM 82/83 ambulance and AMBUS |

(Note to Offerors: Specific performance values and detailed test criteria are contained in the Attachments to the Statement of Objectives.)

SECTION H: POINT OF CONTACT

Agreements Officer: Dan L. Adams, Jr.

Email: dan.l.adams2.civ@army.mil

Agreements Officer Agreements Specialist: Christine Halamoutis
Email: christine.m.halamoutis.civ@army.mil