1.0 Scope of Work:

The Histology Department of Pathology & Laboratory Medicine (P&LMS) is requesting Tissue Processors for use in the Histology Departments at the Greater Los Angeles VA, Phoenix VA?, San Diego VA, Loma Linda VA, Long Beach VA, Tucson VA and New Mexico VA.  Also required is a service and maintenance agreement to start immediately after the warranty period ends. The proposed use of the processors is to process tissue specimens received from the operating room and clinics.

The supply acquisition does not involve connection of contractor-owned IT devices to the VA internal network, therefore the C&A requirements do not apply and a security accreditation package is not required.  The instrument purchased does not have the ability to store data or accept patient demographic information.

2.0 Tasks:  STANDARD

The Contractor shall provide all products, labor, and materials according to the SOW Specifications.  The items required must meet or exceed the specifications listed below.

General Requirements; the contractor shall provide the following: STANDARD

- Specifications set forth in this proposal are minimum specifications and shall not be construed as limiting the overall quality, quantity, or performance of the equipment.

- The contractor shall meet or exceed the minimum requirements and shall be held responsible for the performance and overall quality of the equipment during the warranty period.

- The equipment, based upon the specifications requested, shall be the newest and the most current model.  A refurbished system is not acceptable.  The equipment should be a standard product of the manufacturer of record.  The manufacturer of record is defined as a company whose main occupation is the manufacture for sale items of equipment and which:

- Maintains a factory production line.

- Maintains a stock of replacement parts for the item.

- Maintains engineering drawings, specifications, operating manuals and maintenance manuals.

- Has published and distributed descriptive literature and equipment specifications on the equipment.

2.1 Salient Characteristics; the contractor shall provide the following:

The requested equipment shall have the following salient characteristics:

Tissue Processor

- Instrument Software – Internal to testing instrumentation that provides/maintains the   following: Ability to monitor instrument performance.

- Provide conventional processing of human tissue for clinical and research histology

- Continuous monitoring of vital functions with immediate operator notification of failure(s).

- The tissue processor should manage reagent rotation.

- Ability to use non-proprietary reagents

- The tissue processor/s must accommodate a total of 600 cassettes using 1 or more instruments, with a minimum capacity of 300 cassettes per run, e.g.

-  2 instruments processing 300 specimens each or

-  1 instrument capable of processing 600 specimens.

- Proposal must include potential volume capacity for individual runs on the instrument for the following:

  - Small biopsy

  - Large tissue

  - Fatty tissue

- Automated reagent reconstitution

- Monitor reagent volumes during operation

- Automatic solution transfer

- Temperature controlled retort

- The tissue processor should allow for xylene-free processing.

- Closed system

- Dual retorts are preferred

- Solution detection and management

- Ability to have delayed start times

- Ability to program multiple protocols

- Filtering system to reduce noxious fumes

- Any additional costs such as industrial funding fees, special shipping and handling fees, or other usage fees must be clearly indicated in the price quote.

- Any/all additional necessary accessories required for operating the tissue processor should be listed as an addendum in the pricing schedule.

- Power requirement: Single phase, 115 VAC ± 10%, 50/60 Hz, 12 A

- Include Paraffin and Reagent resivoirs

- Level sensors: at least 4

Uninterruptible Power Supply (UPS)

- One (1) UPS shall be included for each instrument proposed and have

- Line conditioner (be line-interactive)

- Provide electrical power for a minimum of 20 minutes after power fails

- Must allow for an automatic controlled shutdown to prevent damage to the instrument and data records

- Audible alarm when on battery during power fails,  low battery alarm and Overload alarm

- Display showing battery status and charge

- At least 4 power outlets

- Output Operation Voltage (Max): 120 V

2.5 Pre-Installation Assessment; the contractor shall perform the following: STANDARD

- The vendor/distributor must provide all necessary procedure manuals, troubleshooting manuals, and operator manuals.

- The vendor will provide environmental requirements (heat/cool, humidity, light, etc.) and recommendations for instrument and testing material storage.

- The manufacturer of the instrument system should define maintenance requirements and indicate what maintenance is performed by the end-user and what maintenance is performed by the manufacturer.

- Technical capability:

- Provide dimensions of the system (key element)

- List utility requirements (electrical, heat/cool, humidity, etc.)

- Describe maintenance requirements for daily/weekly/monthly maintenance.

- Describe any environmental impact issues that instrumentation may create (water requirements, media disposal, etc.).

- Describe user friendliness:

- Describe ability to trouble-shoot system.

2.6 Support Services: The contractor shall perform the following:  ALTER AS NECESSARY

- The equipment manufacturer should have the ability to provide service sufficient to meet the needs of the user facility Monday-Friday 8:00am to 4:30pm local time.

- Technical support should be available Monday-Friday 8:00am to 4:30pm local time and should be through the manufacturer of the instrument.

3.0 Mandatory Value-Added Characteristics:  STANDARD

These characteristics are mandatory, and will be considered in the best-value determination.  Value added characteristics include, but are not limited to:

- Provide utility requirements (electrical, media disposal, etc.) which should include all necessary SDS, OSHA and safety documents.

- Provide environmental requirements (heat/cool, humidity, light, etc.) and recommendations for [instrument] storage.

- The manufacturer should define maintenance requirements and indicate what maintenance is performed by the user and what is not.

4.0 Expected Period of performance

- 90 days After Receipt of Order (ARO), all equipment must be delivered and installed?

Required Items List:

| Brand Name | MFR part # | Description | QTY (up to) |
| --- | --- | --- | --- |
| Sakura Finetek | 6040 | Tissue Tek VIP 6 AI Vacuum | 14 |
| Sakura Finetek | 6033 | Tissue Tek Uninterrupted Power Supply (UPS), 30 Minute | 14 |

This is a brand name or equal RFI.  The equal products must be able to provide the brand name function or better per the salient characteristics in section 2.1 herein.

The quantities are being finalized at the time of this posting and will depend on the final capacity of the accepted instrument if a solicitation is issued.