# DRAFT PRODUCT DESCRIPTION

# DEPARTMENT OF VETERANS AFFAIRS

VETERANS HEALTH ADMINISTRATION

BIRMINGHAM MEDICAL CENTER

Brand Name or Equal Hillrom Voalte Nurse Call System

Date: April 9, 2026

PD Version Number: .01

# Contents

1.0	PRODUCT REQUIREMENTS	3

1.1	SALIENT CHARACTERISTICS	23

1.2	PROJECT COORDINATION	24

1.3	INSTALLATION AND CONFIGURATION	26

1.4	TRAINING	27

1.5	WARRANTY, MAINTENANCE AND TECHNICAL SUPPORT	28

2.0	NOTICE OF THE FEDERAL ACCESSIBILITY LAW AFFECTING ALL INFORMATION AND COMMUNICATIONUNICATION TECHNOLOGY (ICT) PROCUREMENTS (SECTION 508)	28

2.1	SECTION 508 – INFORMATION AND COMMUNICATIONUNICATION TECHNOLOGY (ICT) STANDARDS	29

2.2	COMPATABILITY WITH ASSISTIVE TECHNOLOGY	29

2.3	ACCEPTANCE AND ACCEPTANCE TESTING	29

3.0	SHIPMENT OF HARDWARE OR EQUIPMENT	30

3.1	SPECIAL SHIPPING INSTRUCTIONS	30

4.0	GENERAL REQUIREMENTS	31

4.1	VA TECHNICAL REFERENCE MODEL	31

4.2	ZERO TRUST – VA CRITICAL SECURITY CONTROLS	31

4.3	SOCIAL SECURITY NUMBER (SSN) REDUCTION	31

4.4	INTERNET PROTOCOL VERSION 6 (IPV6)	32

4.5	SOFTWARE AND LICENSING REQUIREMENTS	32

4.6	TRUSTED INTERNET CONNECTION (TIC)	32

4.7	GENERATIVE ARTIFICIAL INTELLIGENCE REQUIREMENTS	32

5.0	SECURITY AND PRIVACY REQUIREMENTS	34

5.1	POSITION/TASK RISK DESIGNATION LEVEL(S)	34

5.2	CONTRACTOR PERSONNEL SECURITY REQUIREMENTS	35

ADDENDUM B – VA INFORMATION AND INFORMATION SYSTEM SECURITY/ PRIVACY LANGUAGE	37

# PRODUCT REQUIREMENTS

The Contractor shall provide the Department of Veterans Affairs (VA) Veterans Health Administration (VHA) Birmingham VA Health Care System (BVHCS) with a brand name or equal Hillrom Voalte nurse call system with code blue communication system, staging services, installation services, termination/testing/labeling of cables, warranty, and installation. The Contractor shall also provide server/switch implementation in the BVHCS main server room. The system shall integrate with the 107 Hillrom beds that are in use at the BVHCS. The contractor shall provide an installation plan for implementation of the new system and removal of the existing system.

Upon completion of installation, the Contractor shall provide an inventory of what was installed to include equipment type, location, manufacturer, model, and serial number.

The Contractor shall provide the following brand name or equal Hillrom Voalte products/services for each of the BVHCS locations:

| Item # | Product Description | Part | Quantity |
| --- | --- | --- | --- |
| 0001 | Ground Floor Positron Emission tomography Computed tomography |  |  |
|  | Graphical Room Station (GRS) 10H Desk Mount, 4 Universal Serial Bus (USB) | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 5 |
|  | Voalte Nursecall Graphical Room Station 7 Revision 3 without Code | P2594VNC7D00 | 1 |
|  | Equipment Receptacle (Trade Agreements Act (TAA) Compliant) | P2516A01-G-1 | 5 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 5 |
|  | Dome Light, White – TAA Compliant | P2506NNC8A02-G-2 | 11 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 1 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 1 |
|  | B Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 5 |
|  | B Adapter Plate – Nurse Communications Module to Navicare Nurse Call (NNC) | P2523C27-2 | 9 |
|  | B Room Control Board (RCB) 3 Wo Enclosed-Enhanced (Encl-Enhced), Remote Control Box Component (UPOE Cmpt) | P2599NNC3D00 | 7 |
|  | B Enclosure, Room Box, Assembly (Assy) | P2549A03-4 | 7 |
|  | B Communication Cable, 4 Pair (Pr), Category 6, Plenum | P2536A0701-2 | 4 |
|  | B RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | B POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | Uninterruptible Power Supply (UPS) Rack Mount (Rkmt) 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 7 |
| 0002 | Main 1st Dialysis |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 3 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 11 |
|  | Voalte Nursecall Graphical Room Station 7 Revision 3 without Code | P2594VNC7D00 | 1 |
|  | Cable Assembly, Light/Television Control 18 | P2535A0101-4 | 11 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 11 |
|  | Pill Speaker, Digital Audio (Da), Microphonerophone, Device (Dv), Nurse (N), 2 Light (Lt), 8 Feet | P25104C331U-Z105018-2 | 11 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 13 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 2 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 12 |
|  | RCB 3 without Encl-Enhced, Remote Control Box Component (UPOE Cmpt) | P2599NNC3D00 | 10 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 10 |
|  | Communication Cable, 4 Pair (Pr), Category 6, Plenum | P2536A0701-2 | 5 |
|  | Power Over Ethernet (POE) Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 10 |
| 0003 | Main 1st Floor Emergency Department |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 2 |
|  | System Reporting Service without Code, 3 Button, Revision B | P2594NNC1B01-6 | 4 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Equipment Receptacle (TAA Compliant) | P2516A01-G-1 | 20 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 20 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 30 |
|  | BH Push Button Switch, with Gasket | P2520BHA00 | 6 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 1 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 1 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 5 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 1 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 18 |
|  | BH System Reporting Service Cover without Code 3 Button | P2594NNC1B31-1 | 4 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 18 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 18 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 9 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 3 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 2 |
|  | NNC Features-Smart Client | P25NNC007-4 | 2 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 18 |
| 0004 | Main 1st Oncology |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 4 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 1 |
|  | Equipment Receptacle (TAA Compliant) | P2516A01-G-1 | 11 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 11 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 26 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 10 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 10 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 5 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 24 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 17 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 17 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 17 |
| 0005 | Radiology |  |  |
|  | GRS 10H Wall Mount without Handset | P2594VNC5C01 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | Voalte Nursecall GRS 7 Revision 3 without Code | P2594VNC7D00 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 47 |
|  | BH Push Button Switch, with Gasket | P2520BHA00 | 4 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 24 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 22 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 13 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 38 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 32 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 32 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 13 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 5 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 4 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 5 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 3 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 32 |
| 0006 | 4th Floor 4 Main B |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 3 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 18 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 18 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 18 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 18 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 2 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 15 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 15 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 16 |
|  | Adapter Plate, 4G Blank | P2523C3604-1 | 2 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 12 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 12 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1440Va Non-Seismic | P2521B02-1 | 1 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 12 |
| 0007 | 4th Floor 4A |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 3 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 18 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 18 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 18 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 18 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 2 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 15 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 15 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 16 |
|  | Adapter Plate, 4G Blank | P2523C3604-1 | 2 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 12 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 12 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RH45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 2 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 12 |
| 0008 | 4th Floor Surgical Intensive Care Unit |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 3 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 9 |
|  | Voalte Nc Graphical Room Station 7 Revision 3 without Code | P2594VNC7D00 | 1 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 7 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 7 |
|  | Equipment Receptacle (TAA) | P2516A01-G-1 | 4 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 7 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 4 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 17 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 2 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 12 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 14 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 14 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 6 |
|  | Rj45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 3 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 3 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 14 |
| 0009 | 6th Floor Pulmanary |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 6 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 2 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 2 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 3 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 5 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 5 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 5 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 3 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 1 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 5 |
| 0010 | 6th Floor 6B |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 5 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 13 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 13 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 13 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 13 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 14 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 13 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 13 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 13 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 12 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 12 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 7 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 12 |
| 0011 | 6th Floor Cardiology B |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 12 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 9 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 8 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 1 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 9 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 9 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 4 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 9 |
| 0012 | 6th Floor Critical Care Unit |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 4 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 8 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 8 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 8 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 8 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 11 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 9 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 9 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 9 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 9 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 5 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 9 |
| 0013 | 6th Floor Medical Intensive Care Unit |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 2 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 10 |
|  | Voalte Nc Graphical Room Station 7 Rev3 without Code | P2594VNC7D00 | 2 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 10 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 10 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 10 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 11 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 1 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 11 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 10 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 10 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 5 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 10 |
| 0014 | 6th Floor Safe Harbor |  |  |
|  | GRS 10H Desk Mount, 4USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 4 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 10 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 10 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 10 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 10 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 16 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 13 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 15 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 14 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 12 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 12 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 7 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 12 |
| 0015 | 4TH Floor Recovery |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 2 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 2 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 2 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 2 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 6 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 1 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 6 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 9 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 10 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 6 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 6 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 3 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 6 |
| 0016 | 4TH Floor Operating Room |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code, 3 Button, Rev B | P2594NNC1B01-6 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 11 |
|  | Equipment Receptacle (TAA) | P2516A01-G-1 | 6 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 6 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 11 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 6 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 5 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 17 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 17 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 7 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1500VA TAA Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 17 |
| 0017 | 3rd Floor Server Room |  |  |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 1 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 1 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 1 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 1 |
| 0018 | 4TH Floor Recovery |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 2 |
|  | System Reporting Service without Code, 3 Button, Rev B | P2594NNC1B01-6 | 4 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Equipment Receptacle (TAA) | P2516A01-G-1 | 20 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 20 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 30 |
|  | BH Push Button Switch, with Gasket | P2520BHA00 | 6 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 18 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 18 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 18 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 9 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 3 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 2 |
|  | NNC Features-Smart Client | P25NNC007-4 | 2 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 18 |
| 0019 | Main 3rd Floor 3KCLINIC |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | Voalte Nc Graphical Room Station 7 Rev3 without Code | P2594VNC7D00 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 12 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 3 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 3 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 4 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 8 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 10 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 10 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 4 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 10 |
| 0020 | Main 3rd Floor Blind Rehabilitation Clinic Education |  |  |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 1 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 1 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 1 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 1 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 1 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 2 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 2 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 1 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 1 |
|  | Graphical Room Station 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 2 |
| 0021 | Main 3rd Floor Switchboard Operator |  |  |
|  | GRS 10H Wall Mount without Handset | P2594VNC5C01 | 2 |
|  | RCB 3 Wo Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 2 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 2 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 2 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 1 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 2 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 2 |
| 0022 | Main 3rd Floor Ambulatory Services |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 7 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 12 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 12 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 12 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 12 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 31 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 9 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 7 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 15 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 22 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 22 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 9 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 2 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 22 |
| 0023 | Blind Rehabilitation Clinic 4th Floor |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 13 |
|  | Equipment Receptacle (TAA Compliant) | P2516A01-G-1 | 13 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 13 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 20 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 1 |
|  | Push Button Switch, Staff Emergency, with Cancel | P2520NNC3A08-2 | 1 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 8 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 10 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 17 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 13 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 13 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 7 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 3 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1500VA TAA Compliant Seismic | P2521A02-G | 1 |
|  | NNC Features- Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 13 |
| 0024 | Blind Rehabilitation Clinic 5th Floor |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 15 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 15 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 15 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 22 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 11 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 9 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 16 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 16 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1500VA TAA Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 16 |
| 0025 | Main 5th Floor 5A |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 2 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 18 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 18 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 18 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 23 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 6 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 17 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 16 |
|  | Adapter Plate, 3G to 1G | P2523C19-3 | 6 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 13 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 13 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 5 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 13 |
| 0026 | Main 5th Floor 5B |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 2 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 15 |
|  | Cable Assy, Light/Television Control 18 | P2535A0101-4 | 16 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 16 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 16 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 20 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 2 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 15 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 15 |
|  | Adapter Plate, 3G to 1G | P2523C19-3 | 2 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 13 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 13 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-1 | 1 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 8 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 4 |
|  | POE Switch NNC 24-Port UJ60950 | P2519NNC2A24-3 | 2 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1500VA TAA Compliant Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 13 |
| 0027 | 6th Floor Blind Rehabilitation Clinic |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | Voalte Nursecall Graphical Room Station 7 Rev3 without Code | P2594VNC7D00 | 1 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 8 |
|  | Push Button Switch, Code Blue, with Cancel | P2520NNC3A07-2 | 5 |
|  | Push Button Switch, Staff Emergencygency, with Cancel | P2520NNC3A08-2 | 5 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 2 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 7 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 6 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 6 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 3 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | UPS Rkmt 2U 1500VA TAA Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 6 |
| 0028 | 7th Floor Sleep Lab |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 6 |
|  | Equipment Receptacle (TAA Compliant) | P2516A01-G-1 | 6 |
|  | Call Cord, Oxygen Safe, 8 Feet | P2511SCC-0101-080-U-2 | 6 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 8 |
|  | Pull Switch, Bath withcancel | P2520NNC3B01-2 | 7 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 6 |
|  | Adapter Plate - Nurse Communications Module to NNC | P2523C27-2 | 7 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 5 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 5 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 3 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 2 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1500VA TAA Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 5 |
| 0029 | 8th Floor |  |  |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 2 |
|  | Pull Switch, Bath with cancel | P2520NNC3B01-2 | 5 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 2 |
|  | Enclosure, Room Box, Assy | P2549A03-4 | 2 |
|  | Communication Cable, 4Pr, Category 6, Plenum | P2536A0701-2 | 2 |
|  | RJ45 Plug, 100 Pack, Category 6 & 5E | P2537C0101-3 | 1 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | UPS Rkmt 2U 1500VA TAA Seismic | P2521A02-G | 1 |
|  | NNC Features-Smart Client | P25NNC007-4 | 1 |
|  | NNC Enhanced Base Software | P25NNC096-2 | 2 |
| 0030 | Enterprise Servers |  |  |
|  | Aruba 5406R ZL2 CDW4361359 | P009665-1 | 1 |
|  | Aruba 5400R 1100W CDW4361365 | P009667-1 | 2 |
|  | Aruba 24P 1GBE CDW4361377 | P009663-1 | 2 |
|  | Aruba POE+ 12P 1GBE CDW4361378 | P009666-1 | 1 |
|  | Aruba 1G SFP LC CDW4921521 | P009658-1 | 26 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 1 |
|  | Fiber Optic Small Form-Factor Gigabit Interface Converter | P2519NNCFA00-2 | 2 |
|  | Implementation Services | P2570IMPLSERV-3 | 16 |
|  | Project Management Services | P2570PMSERVICE-5 | 2 |
|  | Voice Over Internet Protocol Server By Others (Veterans Affairs (VA)) | SERVER-VOIP | 1 |
|  | Enterprise Server By Others (VA) | SERVER-ENTERPRISE | 1 |
|  | Database Server By Others (VA) | SERVER-DATABASE | 1 |
|  | Application Server By Others (VA) | SERVER-APPLICATION | 1 |
|  | Structured Query Language Server Licenses By Others (VA) | SQL-SERVER-LICENSES | 1 |
| 0031 | Enterprise Installation |  |  |
|  | Installation By Hillrom | P2570A0003-BRM-V | 1 |
| 0032 | Enterprise Implementation Services |  |  |
|  | Implementation Services | P2570IMPLSERV-4 | 578 |
| 0033 | Enterprise Project Management Services |  |  |
|  | Project Management Services | P2570PMSERVICE-5 | 223 |
| 0034 | Enterprise Clinical Services |  |  |
|  | Clinical Services | P25NNC411-HRLY | 149 |
| 0035 | Enterprise Reporting |  |  |
|  | HR Enterprise Reporting License-1 Per Unit | P25NNC072-3 | 30 |
|  | HR Reporting- User License-Per Nursing Unit | P25NNC075-3 | 30 |
|  | HR Reporting Services (Time) | P25NNC066-2 | 1 |
|  | Clinical Consulting Day | P25NNC411-2 | 4 |
|  | HR Reporting - (One-Year Service Maintenance Agreement) | P25NNC078-BRM | 1 |
| 0036 | Enterprise Wireless |  |  |
|  | Wireless Facility Software License | PS102WI-FACILITY-3 | 1 |
|  | Wireless Facility Services | PS102WI-FACILITY-SVCS-3 | 1 |
| 0037 | Enterprise Staff Assignment |  |  |
|  | Staff Assignment Facility Software | PS020SAI-VA | 1 |
|  | Staff Assignment Facility Services | PS020SAISVCS-1 | 1 |
| 0038 | Spare Parts |  |  |
|  | GRS 10H Desk Mount, 4 USB | P2500VNC3D00 | 2 |
|  | Equipment Receptacle (TAA Compliant) | P2516A01-G-1 | 5 |
|  | Dome Light, White -TAA Compliant | P2506NNC8A02-G-2 | 3 |
|  | System Reporting Service without Code Voalte Nursecall | P2594VNC1F01 | 5 |
|  | System Reporting Service with Code Voalte Nursecall | P2594VNC1F11 | 10 |
|  | Audio Station Bed Connector | P2505NNC1C00-1 | 10 |
|  | Pill Speaker, Da, Microphone, Dv, N, 2 Lt, 8 Ft | P25104C331U-Z105018-2 | 10 |
|  | Pull Switch, Bath with Cancel | P2520NNC3B01-2 | 2 |
|  | Pull Switch, Bath, without Cancel | P2520NNC3B02-2 | 2 |
|  | RCB 3 without Encl-Enhced, UPOE Cmpt | P2599NNC3D00 | 7 |
|  | POE Switch NNC 24-Port UL60950 | P2519NNC2A24-3 | 2 |
| 0039 | Technical On-Site Training |  |  |
|  | 3-Day Customer Host Technical Training | P25NNC01503-1 | 1 |
| 0040 | Enterprise Admission, Discharge, Transfer (ADT) |  |  |
|  | ADT Facility Software License | P25NNC008-2 | 1 |
|  | ADT Facility Services | PS023ADT-2 | 1 |
|  | UPS Rkmt 2U 1440VA Non-Seismic | P2521B02-2 | 1 |

## SALIENT CHARACTERISTICS

The following are salient characteristics of a brand name or equal Hillrom Voalte nurse call system:

The brand name or equal Hillrom Voalte nurse call system shall not use middleware to interface bed data.

A brand name or equal Hillrom Voalte nurse call system shall be approved on the Medical Device Risk Assessment (MDERA).

A brand name or equal Hillrom Voalte nurse call systems with code blue communication system shall provide all information required below in a visual dashboard accessed via the workstations located at the nurse stations. The dashboard shall list the following required information:

- Patient room number

- Patient name or corresponding Health Insurance Portability and Accountability Act (HIPPA) compliant naming convention.

- Patient notes (ex. Hard of hearing, NPO (nothing by mouth), etc.)

- Risk Status: based on electronic medical records assessment (Fall Risk, Skin, Pulmonary) or manual entry

- Assigned staff

- Active calls

- Call wait times

- Rounding reminders that automatically reset based on caregiver presence

- Caregiver presence displays on Status Board, Dome Light illumination to signal staff for patient assistance according to the hospital’s pre-determined code comprised of lights and sounds, graphical stations

- Patient presence in bed

- Bed exit armed

- Brakes set

- Bed in low position

- Individual Side rails status, i.e., up/down

- Head of bed angle

- Patient weight

The brand name or equal Hillrom Voalte Nurse Call system shall include Room Control Boards (RCB) that can function in a standalone mode while experiencing loss-of-system communication and homerun or equivalent wiring topology.

## PROJECT COORDINATION

The Contractor shall provide a Contractor Project Manager (CPM) who shall act as a liaison between the Contractor and the Government. The CPM shall be authorized by the Contractor to represent the Contractor in all activities.

The CPM shall ensure that the Contractor deliverables as set forth in this document meet the requirements of the PD and are delivered and installed. The CPM shall schedule ongoing meetings with the Government POCs as listed in Section 4 of this PD as needed throughout this effort. The CPM shall coordinate and hold a kick-off meeting for the effort no later than 60 days after award. The Contractor shall work with the Government POCs listed in Section 4 of this PD to ensure that the following are invited to the kickoff meeting: Delivery Date Coordinator, Implementation Manager, and Facility CIOs (or designee) to discuss delivery schedule requirements and facilitate delivery of equipment.

At the kick-off meeting, the Contractor shall provide a draft implementation project plan to the Government POC as listed in Section 4 of this PD no later than 30 days after award. The document will be used at the kick-off meeting as a basis for discussion. The CPM shall work with the Government to assure a common understanding of work to be performed and responsibilities. The Contractor shall maintain the implementation plan and update it based on environmental changes. The Government will review and approve the implementation plan. At the kick-off meeting the Government Delivery Date Coordinator, Implementation Manager, and Facility CIOs (or designee) will discuss delivery schedule requirements and facilitate delivery of equipment.

At a minimum, the implementation plan shall include:

- A list of Government and Contractor POCs and roles.

- System Documentation and low-level system design and high-level system design documentation. An initial draft shall be presented at the kick-off meeting to the Government POCs listed in PD Section 4 for review and approval.

- Project Management Schedule. The Implementation plan shall identify specific deployment tasks, milestones, and timelines.

  - The project management schedule shall identify significant milestones to accomplish the requirements of this PWS. The requirements, including delivery schedule, shall be listed by conference room.

- Master Delivery Schedule. The Master Delivery Schedule shall identify all products/items to be delivered and the status. All products/items shipped shall include the date shipped and the serial number associated with each product/item, and the anticipated date for the Government to receive the product/item. The Master Delivery Schedule shall reference Shipment Reports. Shipment reports shall be submitted to the Government POCs listed in Section 4 of this PD. Shipment Reports shall be in Microsoft Excel Format and shall clearly identify each serial number of the equipment being delivered, i.e., only one serial number per cell.

- The contractor shall provide and maintain an accurate detailed Inventory List of the Contractor installed equipment and firmware including serial numbers, Equipment Entry (EE) Tag numbers, version release, and licenses.

- Risks or Issues: The implementation plan shall identify and document all risks or issues that are anticipated or arise during performance and recommend strategies to mitigate the identified risks.

- The contractor shall provide a Final Report to include a compilation all installation reports.

- The contractor shall provide Electronic PDF copy of “As built” drawings.

- The contractor shall provide a copy of final uncompiled Crestron Code and audio processor file.

The CPM shall work with the Government POCs identified in Section 4 of this PD to ensure that all installations are scheduled and performed with minimal interruption to patient care. The VA will make every effort to schedule installation during normal business hours, however there may be instances where night and weekend hours may be necessary to complete the installation to ensure that patient care is minimally impacted.

## INSTALLATION AND CONFIGURATION

The Contractor shall remove and replace all components of the nurse call System including but not limited to the headend amplifier system, speakers, components, conductors, conduit, etc. located throughout the facilities.

The Contractor shall provide all labor, materials, equipment, etc. necessary to replace the Nurse Call with Code Blue System in accordance with the PD, the contract documents, National Fire Protection Agency (NFPA), local building codes, Underwriters Laboratories (UL) 1069 Standard for safety for hospital signaling and nurse call equipment, and VA and American Society of Healthcare Engineering (ASHE) facility standards.

Wiring infrastructure shall be installed by the contractor. The contractor shall then work with the Government to test system functionality prior to activating the rooms/areas. The Contractor shall obtain written notice from the Government’s Contracting Officer’s Representative (COR) to proceed with activation of the room(s)/areas. Activation of room(s)/areas should only impact operations with a partial-day outage. The Contractor shall obtain written acceptance from the COR for each area/ward after activation.

In accordance with VA Handbook 6513, the brand name or equal Hillrom Voalte Nurse Call system shall have a VA approved Memorandum of Understand/Interconnection Security Agreement (MOU/ISA) E-5326 in place with VA that allows remote access to the Virtual Servers planned to be installed.

All materials and/or equipment necessary for complete and proper operation of the brand name or equal Hillrom Voalte Nurse Call system, even if not specifically mentioned in this document, shall be deemed part of the contract.

All equipment shall be installed and connected in strict accordance with the Original Equipment Manufacturer’s (OEM) recommended instructions.

The Contractor shall ensure all hallways and walkways are clear of construction materials and debris upon departure, including packing materials.

The Contractor shall be responsible for disposal of all packing materials and shall remove them from the facility. Dumpsters on campus may be used with prior approval to dispose of debris and packaging materials.

The Contractor shall adhere to standards of good cable management, including bundling and labeling all cables.

The contractor shall provide installation of all products for the brand name or equal Hillrom Voalte Nurse Call system with code blue system to function as intended.

The Contractor shall travel to the installation sites, as listed in Section 1.X of this PD, to complete the tasks in this PD. Any necessary travel shall be included in the installation price.

The contractor shall connect the brand name or equal Nurse Call system to the VA network on Medical Device-Legacy Information Technology Environment (MD-LITE).

The Contractor shall deliver and install all components of the brand name or equal Nurse Call system and complete connections to the servers provided for this system. The Contractor shall configure all hardware and software, update any embedded firmware required for activation or network connection. The Contractor shall test/inspect the system after installation to ensure the brand name or equal Nurse Call system functions as intended and ensure proper operation. The Contractor shall ensure the installation and configuration align with the Original Equipment Manufacturer specifications. The Contractor shall provide the manufacturer’s standard warranty for the system installed.

No work at the Government site shall take place on Federal holidays, weekends, or outside of normal business hours unless directed by the Contracting Officer (CO).

Tasks under this PD shall be performed at the following VA facility:

Birmingham VA Healthcare System

Address:	Birmingham VA Medical Center

700 19th Street South

Birmingham, AL 35233

## TRAINING

The Contractor shall coordinate and provide training sessions to ensure VA staff have the product knowledge and skills necessary to both use and keep the brand name or equal Nurse Call system functioning as intended to ensure reliable communication and patient safety.

The Contractor shall provide on-site clinical training for end users and technical training for HTM/Biomed staff who are responsible for maintaining the system.

Clinical training shall use a “train the trainer” approach and include the following:

  - Product overview

  - Alarm prioritization

  - Patient education

  - Call types

  - Integrated system functionality

  - Workflow and protocols

  - Reporting and data analytics

  - How to operate the system.

Technical training shall include the following:

  - Product overview

  - Specific equipment installed

  - General system operation

  - How to troubleshoot common issues

  - How to perform required maintenance

The Contractor shall provide quick reference sheets for the training sessions that the Government can use in the future to train new clinical staff.

## WARRANTY, MAINTENANCE AND TECHNICAL SUPPORT

The Contractor shall provide the brand name or equal Hillrom Voalte manufacturer’s warranty at no additional cost to the Government for a period of 12 months, commencing upon Government acceptance of all products/services herein. The Contractor shall ensure all installation and services provided comply with the OEM’s warranty requirements.

The Contractor shall be certified by the OEM to perform all installation, configuration, warranty, and maintenance and technical support services.

The warranty shall include maintenance and technical support for the brand name or equal Nurse Call system 24 hours a day x 7 days a week x 365 days a year, including:

- Unlimited phone support.

- On-site service if the contractor cannot resolve the issue during phone support.

Maintenance and technical support shall include:

- All firmware major and minor releases, patches, bug fixes, to ensure the firmware, firmware (if included) and equipment function as intended.

- Rapid telephone support response time of no longer than two hours

- On-site response time within 24 hours

- Help Desk support

- Expedited freight if replacement is required

- All parts, equipment, labor, and travel required to ensure the brand name or equal Nurse Call system function as intended

# NOTICE OF THE FEDERAL ACCESSIBILITY LAW AFFECTING ALL INFORMATION AND COMMUNICATIONUNICATION TECHNOLOGY (ICT) PROCUREMENTS (SECTION 508)

On January 18, 2017, the Access Board issued a final rule that updated accessibility requirements covered by Section 508 and refreshed guidelines for telecommunications equipment subject to Section 255 of the Communications Act. The final rule went into effect on January 18, 2018. The revisions and updates to the Section 508-based standards and Section 255-based guidelines are intended to ensure that information and communication technology (ICT) covered by the respective statutes is accessible to and usable by individuals with disabilities.

## SECTION 508 – INFORMATION AND COMMUNICATIONUNICATION TECHNOLOGY (ICT) STANDARDS

The Section 508 standards established by the Access Board are incorporated into, and made part of all VA orders, solicitations and purchase orders developed to procure ICT.  These standards are found in their entirety at: Revised 508 Standards and 255 Guidelines (access-board.gov). A single PDF file version of the Revised Section 508 Standards and 255 Guidelines will be supplied upon request, or can be obtained from the Access Board website.  Federal agencies must comply with the Rehabilitation Act of 1973, as amended.

The Contractor shall comply with “508 Chapter 2: Scoping Requirements” for all electronic ICT and content delivered under this contract. Specifically, as appropriate for the technology and its functionality, the Contractor shall comply with the technical standards marked here:

☒	E205 Electronic Content – (Accessibility Standard -WCAG 2.0 Level A and AA Guidelines)

☒	E204 Functional Performance Criteria

☒	E206 Hardware Requirements

☒	E207 SOFTWARE Requirements

☒	E208 Support Documentation and Services Requirements

## COMPATABILITY WITH ASSISTIVE TECHNOLOGY

The standards do not require installation of specific accessibility-related SOFTWARE or attachment of an assistive technology device. Section 508 requires that ICT be compatible with such SOFTWARE and devices so that ICT can be accessible to and usable by individuals using assistive technology, including but not limited to screen readers, screen magnifiers, and speech recognition SOFTWARE.

## ACCEPTANCE AND ACCEPTANCE TESTING

Deliverables resulting from this solicitation will be accepted based in part on satisfaction of the Section 508 Chapter 2: Scoping Requirements standards identified above.

The Government reserves the right to test for Section 508 Compliance before delivery.  The Contractor shall be able to demonstrate Section 508 Compliance upon delivery.

# SHIPMENT OF HARDWARE OR EQUIPMENT

Inspection: Destination
Acceptance: Destination
Free on Board (FOB): Destination

Ship To and Mark For: The Government’s Primary Point of Contact (POC) and Alternate POC To Be Provided (TBP) at award.

Government Primary POC:

Name: TBP

Address: TBP

Telephone: TBP

Email: TBP

Government Alternate POC:

Name: TBP

Address: TBP

Telephone: TBP

Email: TBP

## SPECIAL SHIPPING INSTRUCTIONS

Contractors shall coordinate all deliveries with the COR and POCs in section 4.0 of this PD before sending any shipments to ensure timely receipt and adequate storage space will be available.

Prior to shipping, Contractor shall notify the COR and POCs in section 4.0 by email, of all incoming deliveries.

All shipments, either single or multiple container deliveries, shall bear the VA IFCAP Purchase Order number on external shipping labels and associated manifests or packing lists. In the case of multiple container deliveries, a statement readable near the VA Integrated Funds Distribution, Control Point Activity, Accounting and Procurement (IFCAP) Purchase Order (PO) or Contract number shall indicate total number of containers for the complete shipment (e.g. “Package 1 of 2”), clearly readable on manifests and external shipping labels.

Packing Slips/Labels and Lists shall also include the following:

IFCAP PO #: See block #20 of the Standard Form, SF 1449

Project Description: Brand Name or Equal Hillrom Voalte Nurse Call System

Total number of Containers:  Package ___ of ___.  (e.g., Package 1 of 3)

# GENERAL REQUIREMENTS

## VA TECHNICAL REFERENCE MODEL

The Contractor shall comply with the VA OIT Technical Reference Model (VA TRM). Compliance with the VA TRM is achieved by using only technologies and standards that are listed as approved for use in the VA TRM. The Contractor shall provide all necessary information requested by VA to ensure TRM approval is obtained prior to use on VA’s network.

## ZERO TRUST – VA CRITICAL SECURITY CONTROLS

VA has established minimum mandatory security requirements and requires that any network connected SOFTWARE system or service must meet the VA Critical Security Controls as outlined in the VA Memorandum, “VA Security Controls”, https://www.voa.va.gov/DocumentView.aspx?DocumentID=5010.  VA Critical Security Controls identify the minimum mandatory requirements that must be implemented across all VA enterprise infrastructure, cloud computing environments, information systems, networks, and specialized devices (medical devices/systems, special-purpose systems, and research scientific computing devices) that process, store, and/or transmit VA data.  Effective July 1, 2025, the Contractor shall implement these VA Critical Security Controls, within any network connected SOFTWARE system or service prior to being authorized for use in the VA.  This functional requirement is not negotiable, and Plan of Action & Milestones (POAM) will not be accepted in the event these controls cannot be implemented for new systems.  Critical Security Controls are intended to increase VA’s security posture and provide security and privacy risk visibility into the VA network and is not a new requirement.  The Contractor’s failure to maintain these VA Critical Controls after implementation will result in VA discontinuing the use of the system.

## SOCIAL SECURITY NUMBER (SSN) REDUCTION

The Contractor solution shall support the Social Security Number (SSN) Fraud Prevention Act (FPA) of 2017 which prohibits the inclusion of SSNs on any document sent by mail.  The Contractor support shall also be performed in accordance with Section 240 of the Consolidated Appropriations Act (CAA) 2018, enacted March 23, 2018, which mandates VA to discontinue using SSNs to identify individuals in all VA information systems as the Primary Identifier.  The Contractor shall ensure that any new IT solution discontinues the use of SSN as the Primary Identifier to replace the SSN with the Integrated Control Number (ICN) in all VA information systems for all individuals.  The Contractor shall ensure that all Contractor delivered applications and systems integrate with the VA Master Person Index (MPI) for identity traits to include the use of the ICN as the Primary Identifier. The Contractor solution may only use a Social Security Number to identify an individual in an information system if and only if the use of such number is required to obtain information VA requires from an information system that is not under the jurisdiction of VA.

## INTERNET PROTOCOL VERSION 6 (IPV6)

The Contractor solution shall support IPv6-Only based upon the memo issued by the Office of Management and Budget (OMB) on November 19, 2020 (https://www.whitehouse.gov/wp-content/uploads/2020/11/M-21-07.pdf). Which defines IPv6-only as the state of an operational system or service when IPv4 protocol functions (addressing, packet forwarding) are not in use. The NIST USGv6 profile defines technical requirements for a product to be capable of operating in IPv6-Only environments.  IPv6-Only technology, in accordance with the USGv6 Program (https://www.nist.gov/programs-projects/usgv6-program/usgv6-revision-1), NIST Special Publication (SP) 500-267B Revision 1 “USGv6 Profile” (https://doi.org/10.6028/NIST.SP.500-267Br1), and NIST SP 800-119 “Guidelines for the Secure Deployment of IPv6” (https://doi.org/10.6028/NIST.SP.800-119), compliance shall be included in all IT infrastructures, application designs, application development, operational systems and sub-systems, and their integration.  In addition to the above requirements, all devices, applications, and systems shall support all applicable functionality on native IPv6-Only as well as dual stack (IPv6 / IPv4) connectivity without additional memory or other resources being provided by the Government, so that they can function in a mixed environment.  All public/external facing servers and services (e.g., web, email, DNS, ISP services, etc.) shall support native IPv6-Only and dual stack (IPv6 / IPv4) users and all internal infrastructure and applications shall support using native IPv6-Only and dual stack (IPv6 / IPv4) for all functionality and operations.

## SOFTWARE AND LICENSING REQUIREMENTS

The Contractor shall be responsible for the provision of all SOFTWARE licenses and any associated licensing maintenance required for any development, delivery, integration, operation, and/or maintenance associated with its proposed application(s), SOFTWARE products, SOFTWARE solution, and/or system including, but not limited to, any and all application(s), SOFTWARE and/or SOFTWARE products that comprise, are a part of, or integrate with the Contractor’s proposed application(s), SOFTWARE products, SOFTWARE solution, and/or system for the life of any resulting contract.

## TRUSTED INTERNET CONNECTION (TIC)

The Contractor solution shall meet the requirements outlined in Office of Management and Budget Memorandum M-19-26, “Update to the Trusted Internet Connections (TIC) Initiative“ (https://www.whitehouse.gov/wp-content/uploads/2019/09/M-19-26.pdf), VA Directive 6513 “Secure External Connections”, and shall comply with the TIC 3.0 Core Guidance Documents, including all Volumes and TIC Use Cases, found at the Cybersecurity & Infrastructure Security Agency (CISA) (https://www.cisa.gov/publication/tic-30-core-guidance-documents.)

## GENERATIVE ARTIFICIAL INTELLIGENCE REQUIREMENTS

If Artificial Intelligence (AI) is part of the Contractor system/solution, the Contractor shall be in compliance with E.O. 13960 and 14319, OMB Memorandums M-25-21, M-25-22, and M-26-04, and VA Directive and Handbook 6500 to support VA in trustworthy, secure, responsible, accountable, Truth Seeking, and Ideologically Neutral AI innovation.

No web-based, publicly available generative AI service has been approved for use with VA-sensitive data including, but not limited to OpenAI’s ChatGPT, Microsoft’s Copilot, Google’s Gemini, and Anthropic’s Claude.  No PII, protected health information (PHI), or VA-sensitive data should be entered into these unapproved services.  The entire definition of VA Sensitive Data can be found at 38 U.S.C. § 5727 (23).  The Contractor shall protect VA Sensitive data from unauthorized disclosure or use, and from being used to train or improve the functionality of the vendor's commercial offerings without express permission from VA.

The Contractor shall deliver the services, technical data, configurations, models, weights, documentation or other information, including source code, developed during contract performance.  The Contractor, in collaboration with VA, must evaluate the output of any AI tool for accuracy prior to using the output in VA efforts.  All data, including models, weights, algorithms, and outputs generated using either Government-furnished data or data first developed under this contract shall be the property of the Government with unlimited rights as defined by FAR 52.227-14.  The Contractor shall also deliver any custom-developed code developed with AI methodologies in accordance with the SHARE IT Act, if applicable.  VA shall also retain unlimited rights to any improvements to that data, including the continued design, development, testing, and operation of AI systems and/or systems utilizing AI.

The Contractor shall ensure that all AI Systems, including Large Language Models (LLM), whether standalone LLMs or integrated into broader SOFTWARE platforms,  procured under this effort and/or used in performance of this contract complies with the Unbiased AI Principles as defined in Executive Order 14319 and OMB Memorandum M-26-04, e.g. Truth Seeking and Ideologically Neutral, as applicable to the intended use of the system.

The Contractor shall ensure that AI System/LLM generated outputs are truthful in responding to user prompts seeking factual information or analysis, prioritize historical accuracy, scientific inquiry, and objectivity, as well as acknowledge uncertainty where reliable information is incomplete or contradictory and do not fabricate, hallucinate, or present unverifiable information as fact.

The Contractor shall ensure that LLMs and AI outputs are neutral, nonpartisan tools that do not manipulate responses in favor of ideological dogmas such as Diversity, Equity, and Inclusion (DEI).  The Contractor shall ensure that the LLMs, including the AI System, are not intentionally encoded with partisan or ideological judgments within an AI System’s/LLM’s outputs unless those judgments are prompted by or otherwise readily accessible to the end user.

Compliance with the AI requirements of this contract, including adherence to the Unbiased AI Principles, Truth-Seeking and Ideological Neutrality requirements, disclosure obligations, monitoring and reporting requirements, and corrective action responsibilities, is a material condition of eligibility for payment and continued performance. In accordance with OMB Memorandum M-26-04, failure to comply with these requirements, or refusal to take timely and appropriate corrective action to remedy identified noncompliance, may constitute a material breach of contract and may result in remedies consistent with the terms of the contract and applicable law including but not limited to withholding of payment, termination for cause/default, or other appropriate contractual remedies, as determined by the Contracting Officer.

All AI for delivery and/or use in performance of the contract shall adhere to the VA Trustworthy AI Framework, available at https://department.va.gov/ai/trustworthy/.  Once the Contractor has obtained CO authorization to use specific AI in contract performance, the Contractor shall provide the COR a signed Statement of Attestation prior to AI implementation as identified in the Artificial Intelligence in VA Contracts clause, to include its associated list of AI being used in contract performance.

Deliverable:

Statement of Attestation

# SECURITY AND PRIVACY REQUIREMENTS

## POSITION/TASK RISK DESIGNATION LEVEL(S)

In accordance with VA Handbook 0710, Personnel Vetting Program, the position sensitivity and the level of background investigation commensurate with the required level of access for the following tasks within the PD are:

# Position Sensitivity and Background Investigation Requirements

# by Service Task/Item Number

| Service Task / Item Number | Tier1 / Low Risk | Tier 2 / Moderate Risk | Tier 4 / High Risk |
| --- | --- | --- | --- |
| 1.0 | ☒ | ☐ | ☐ |
| 1.1 | ☒ | ☐ | ☐ |
| 1.2 | ☒ | ☐ | ☐ |
| 1.3 | ☒ | ☐ | ☐ |
| 1.4 | ☒ | ☐ | ☐ |
| 1.5 | ☒ | ☐ | ☐ |

The Tasks identified above and the resulting Position Sensitivity and Background Investigation requirements identify, in effect, the Background Investigation requirements for Contractor individuals, based upon the tasks the particular Contractor individual will be working. The submitted Contractor Staff Roster must indicate the required Background Investigation Level for each Contractor individual based upon the tasks the Contractor individual will be working, in accordance with their submitted proposal.

## CONTRACTOR PERSONNEL SECURITY REQUIREMENTS

Contractor Responsibilities:

- The Contractor shall prescreen all personnel requiring access to the computer systems to ensure they maintain the appropriate Background Investigation, and are able to read, write, speak, and understand the English language.

Within 3 business days after award, the Contractor shall provide a roster of Contractor and Subcontractor employees to the COR to begin their background investigations in accordance with the PAL template artifact. The Contractor Staff Roster shall contain the Contractor’s Full Name, Date of Birth, Place of Birth, individual background investigation level requirement (based upon Section 6.2.1 Tasks), etc. The Contractor shall submit full Social Security Numbers either within the Contractor Staff Roster or under separate cover to the COR.  The Contractor Staff Roster shall be updated and provided to VA within 1 day of any changes in employee status, training certification completion status, Background Investigation level status, additions/removal of employees, etc. throughout the Period of Performance. The Contractor Staff Roster shall remain a historical document indicating all past information and the Contractor shall indicate in the comment field, employees no longer supporting this contract. The preferred method to send the Contractor Staff Roster or Social Security Number is by encrypted e-mail. If unable to send encrypted e-mail, other methods which comply with FIPS 140-2 (or its successor) are to encrypt the file, use a secure fax, or use a traceable mail service. (Note that by September 22, 2026, all FIPS 140-2 certificate validations will be placed on the Historical List. See FIPS Transition Effort Timeline for more details:  https://csrc.nist.gov/projects/fips-140-3-transition-effort#schedule)

- The Contractor should coordinate with the location of the nearest VA fingerprinting office through the COR. Only electronic fingerprints are authorized. The Contractor shall bring their completed Personnel Security Adjudication Center (PSAC) Fingerprint request form with them (see paragraph d.4. below) when getting fingerprints taken.

- The Contractor shall ensure the following required forms are submitted to the COR within 5 days after contract award:

- Optional Form 306

- Self-Certification of Continuous Service

- VA Form 0710

- Completed PSAC Fingerprint Request Form

The Contractor personnel shall submit all required information related to their background investigations (completion of the investigation documents (SF85, SF85P, or SF 86) utilizing the Defense Counterintelligence and Security Agency (DCSA) National Background Investigation Services (NBIS) Electronic Application (eAPP) (formerly e-QIP) after receiving an email notification from the PSAC.

The Contractor employee shall certify and release the eAPP document, print, and sign the signature pages, and send them encrypted to the COR for electronic submission to the PSAC. These documents shall be submitted to the COR within 3 business days of receipt of the eAPP) notification email.  (Note:  DCSA is moving towards a “click to sign” process.  If click to sign is used, the Contractor employee should notify the COR within 3 business days that documents were signed via eAPP).

- The Contractor shall be responsible for the actions of all personnel provided to work for VA under this contract.  In the event that damages arise from work performed by Contractor provided personnel, under the auspices of this contract, the Contractor shall be responsible for all resources necessary to remedy the incident.

- A Contractor may be granted unescorted access to VA facilities and/or access to VA Information Technology resources (network and/or protected data) with a favorably adjudicated Special Agreement Check (SAC), completed training delineated in VA Handbook 6500.6 (Appendix C, Section 4), signed “Contractor Rules of Behavior”, and with a valid, operational Personal Identity Verification (PIV) credential for PIV-only logical access to VA’s network.  A PIV card credential can be issued once your SAC has been favorably adjudicated and your background investigation has been scheduled by DCSA.  However, the Contractor will be responsible for the actions of the Contractor personnel they provide to perform work for VA.  The investigative history for Contractor personnel working under this contract must be maintained in the database of DCSA.

- The Contractor, when notified of an unfavorably adjudicated background investigation on a Contractor employee as determined by the Government, shall withdraw the employee from consideration in working under the contract.

Failure to comply with the Contractor personnel security investigative requirements may result in loss of physical and/or logical access to VA facilities and systems by Contractor and Subcontractor employees and/or termination of the contract for default.

PIV Credential Holders must follow all HSPD-12 policies and procedures as well as use and protect their assigned identity credentials in accordance with VA policies and procedures, displaying their badges at all times, and returning the identity credentials upon termination of their relationship with VA.

Deliverable:

- Contractor Staff Roster

# ADDENDUM B – VA INFORMATION AND INFORMATION SYSTEM SECURITY/ PRIVACY LANGUAGE

NOTE: In the event of a conflict, VA Acquisition Regulation (VAAR) Security Clauses take precedence over the language in this Addendum B.

APPLICABLE SECTIONS TAILORED FROM: VA NOTICE 24-12, APRIL 22, 2024, UPDATE TO VA HANDBOOK 6500.6, CONTRACT SECURITY, APPENDIX C VA INFORMATION AND INFORMATION SYSTEM SECURITY/PRIVACY LANGUAGE, FOR INCLUSION INTO CONTRACTS, AS APPROPRIATE

GENERAL

This entire section applies to all acquisitions requiring any Information Security and Privacy language. Contractors, contractor personnel, Subcontractors and subcontractor personnel will be subject to the same federal laws, regulations, standards, VA directives and handbooks, as VA personnel regarding information and information system security and privacy.

VA INFORMATION CUSTODIAL LANGUAGE

- The Government shall receive unlimited rights to data/intellectual property first produced and delivered in the performance of this contract or order (hereinafter “contract”) unless expressly stated otherwise in this contract. This includes all rights to non-proprietary source code and all documentation created in support thereof. The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General.

- Information made available to the Contractor by VA for the performance or administration of this contract will be used only for the purposes specified in the service agreement, SOW, PWS, PD, and/or contract. The Contractor shall not use VA information in any other manner without prior written approval from a VA Contracting Officer (CO). The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General.

- VA information will not be co-mingled with any other data on the Contractor’s information systems or media storage systems. The Contractor shall ensure compliance with Federal and VA requirements related to data protection, data encryption, physical data segregation, logical data segregation, classification requirements and media sanitization.

- VA reserves the right to conduct scheduled or unscheduled audits, assessments, or investigations of Contractor Information Technology (IT) resources to ensure information security is compliant with Federal and VA requirements. The Contractor shall provide all necessary access to records (including electronic and documentary materials related to the contracts and subcontracts) and support (including access to Contractor and Subcontractor staff associated with the contract) to VA, VA's Office Inspector General (OIG), and/or Government Accountability Office (GAO) staff during periodic control assessments, audits, or investigations.

- The Contractor may only use VA information within the terms of the contract and applicable Federal law, regulations, and VA policies. If new Federal information security laws, regulations or VA policies become applicable after execution of the contract, the parties agree to negotiate contract modification and adjustment necessary to implement the new laws, regulations, and/or policies.

- The Contractor shall not make copies of VA information except as specifically authorized and necessary to perform the terms of the contract. If copies are made for restoration purposes, after the restoration is complete, the copies shall be destroyed in accordance with VA Directive 6500, VA Cybersecurity Program and VA Information Security Knowledge Service.

- If a Veterans Health Administration (VHA) contract is terminated for default or cause with a business associate, the related local Business Associate Agreement (BAA) shall also be terminated and actions taken in accordance with VHA Directive 1605.05, Business Associate Agreements. If there is an executed national BAA associated with the contract, VA will determine what actions are appropriate and notify the contractor.

- The Contractor shall store and transmit VA sensitive information in an encrypted form, using VA-approved encryption tools which are, at a minimum, Federal Information Processing Standards (FIPS) 140-2, Security Requirements for Cryptographic Modules (or its successor) validated and in conformance with VA Information Security Knowledge Service requirements. The Contractor shall transmit VA sensitive information using VA approved Transport Layer Security (TLS) configured with FIPS based cipher suites in conformance with National Institute of Standards and Technology (NIST) 800-52, Guidelines for the Selection, Configuration and Use of Transport Layer Security (TLS) Implementations.

- The Contractor’s firewall and web services security controls, as applicable, shall meet or exceed VA’s minimum requirements.

- Except for uses and disclosures of VA information authorized by this contract for performance of the contract, the Contractor may use and disclose VA information only in two situations: (i) in response to a qualifying order of a court of competent jurisdiction after notification to VA CO (ii) with written approval from the VA CO. The Contractor shall refer all requests for, demands for production of or inquiries about, VA information and information systems to the VA CO for response.

- Notwithstanding the provision above, the Contractor shall not release VA records protected by Title 38 U.S.C. § 5705, Confidentiality of medical quality-assurance records and/or Title 38 U.S.C. § 7332, Confidentiality of certain medical records pertaining to drug addiction, sickle cell anemia, alcoholism or alcohol abuse or infection with Human Immunodeficiency Virus (HIV). If the Contractor is in receipt of a court order or other requests for the above-mentioned information, the Contractor shall immediately refer such court order or other requests to the VA CO for response.

- Information made available to the Contractor by VA for the performance or administration of this contract or information developed by the Contractor in performance or administration of the contract will be protected and secured in accordance with VA Directive 6500 and Identity and Access Management (IAM) Security processes specified in the VA Information Security Knowledge Service.

- Any data destruction done on behalf of VA by a Contractor shall be done in accordance with National Archives and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records and Information Management, VA Handbook 6300.1, Records Management Procedures, and applicable VA Records Control Schedules.

- The Contractor shall provide its plan for destruction of all VA data in its possession according to VA Directive 6500 and NIST 800-88, Guidelines for Media Sanitization prior to termination or completion of this contract. If directed by the COR/CO, the Contractor shall return all Federal Records to VA for disposition.

- Any media, such as paper, magnetic tape, magnetic disks, solid state devices or optical discs that is used to store, process, or access VA information that cannot be destroyed shall be returned to VA. The Contractor shall hold the appropriate material until otherwise directed by the Contracting Officer’s Representative (COR) or CO. Items shall be returned securely via VA-approved methods. VA sensitive information must be transmitted utilizing VA-approved encryption tools which are validated under FIPS 140-2 (or its successor) and NIST 800-52. If mailed, the Contractor shall send via a trackable method (USPS, UNINTERRUPTIBLE POWER SUPPLY, FedEx, etc.) and immediately provide the COR/CO with the tracking information. Self-certification by the Contractor that the data destruction requirements above have been met shall be sent to the COR/CO within 30 business days of termination of the contract.

- All electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) used to store, process or access VA information will not be returned to the Contractor at the end of lease, loan, or trade-in. Exceptions to this paragraph will only be granted with the written approval of the VA CO.

ACCESS TO VA INFORMATION AND VA INFORMATION SYSTEMS

- A Contractor/Subcontractor shall request logical (technical) or physical access to VA information and VA information systems for their employees and subcontractors only to the extent necessary to perform the services specified in the solicitation or contract. This includes indirect entities, both affiliate of Contractor/Subcontractor and agent of Contractor/Subcontractor.

- Contractors and subcontractors shall sign the VA Information Security Rule of Behavior (ROB) before access is provided to VA information and information systems (see Section 4, Training, below). The ROB contains the minimum user compliance requirements and does not supersede any policies of VA facilities or other agency components which provide higher levels of protection to VA’s information or information systems. Users who require privileged access shall complete the VA elevated privilege access request processes before privileged access is granted.

- All Contractors and Subcontractors working with VA information are subject to the same security investigative and clearance requirements as those of VA appointees or employees who have access to the same types of information. The level and process of background security investigations for contractors shall be in accordance with VA Directive and Handbook 0710, Personnel Vetting Program. The Office of Human Resources and Administration/Operations, Security and Preparedness (HRA/OSP) is responsible for these policies and procedures. Contract personnel who require access to classified information or information systems shall have an appropriate security clearance. Verification of a Security Clearance shall be processed through the Special Security Officer located in HRA/OSP. Contractors shall conform to all requirements stated in the National Industrial Security Program Operating Manual (NISPOM).

- All Contractors and subcontractors shall comply with conditions specified in VAAR 852.204-71(d); Contractor operations required to be in United States. All Contractors and Subcontractors working with VA information must be permanently located within a jurisdiction subject to the law of the United States or its Territories to the maximum extent feasible. If services are proposed to be performed abroad the Contractor must state where all non-U.S. services are provided. The Contractor shall deliver to VA a detailed plan specifically addressing communications, personnel control, data protection and potential legal issues. The plan shall be approved by the COR/CO in writing prior to access being granted.

- The Contractor shall notify the COR/CO in writing immediately (no later than 24 hours) after personnel separation or occurrence of other causes. Causes may include the following:

  - Contractor/Subcontractor personnel no longer has a need for access to VA information or VA information systems.

  - Contractor/Subcontractor personnel are terminated, suspended, or otherwise has their work on a VA project discontinued for any reason.

  - Contractor believes their own personnel or Subcontractor personnel may pose a threat to their company’s working environment or to any company-owned property. This includes Contractor-owned assets, buildings, confidential data, customers, employees, networks, systems, trade secrets and/or VA data.

  - Any previously undisclosed changes to Contractor/Subcontractor background history are brought to light, including but not limited to changes to background investigation or employee record.

  - Contractor/Subcontractor personnel have their authorization to work in the United States revoked.

  - Agreement by which Contractor provides products and services to VA has either been fulfilled or terminated, such that VA can cut off electronic and/or physical access for Contractor personnel.

- In such cases of contract fulfillment, termination, or other causes; the Contractor shall take the necessary measures to immediately revoke access to VA network, property, information, and information systems (logical and physical) by Contractor/Subcontractor personnel. These measures include (but are not limited to): removing and then securing Personal Identity Verification (PIV) badges and PIV – Interoperable (PIV-I) access badges, VA-issued photo badges, credentials for VA facilities and devices, VA-issued laptops, and authentication tokens. Contractors shall notify the appropriate VA COR/CO immediately to initiate access removal.

- Contractors/Subcontractors who no longer require VA accesses will return VA-issued property to VA. This property includes (but is not limited to): documents, electronic equipment, keys, and parking passes. PIV and PIV-I access badges shall be returned to the nearest VA PIV Badge Issuance Office. Once they have had access to VA information, information systems, networks and VA property in their possessions removed, Contractors shall notify the appropriate VA COR/CO.

TRAINING

- All contractors and subcontractors requiring access to VA information and VA information systems shall successfully complete the following before being granted access to VA information and its systems:

  - VA Privacy and Information Security Awareness and Rules of Behavior course (Talent Management System (TMS) #10176) initially and annually thereafter.

  - Sign and acknowledge (electronically through TMS #10176) understanding of and responsibilities for compliance with the Organizational Rules of Behavior, relating to access to VA information and information systems initially and annually thereafter; and

  - Successfully complete any additional cyber security or privacy training, as required for VA personnel with equivalent information system or information access [to be defined by the VA program official and provided to the VA CO for inclusion in the solicitation document – i.e., any role-based information security training].

- The Contractor shall provide to the COR/CO a copy of the training certificates and certification of signing the Organizational Rules of Behavior for each applicable employee within five days of the initiation of the contract and annually thereafter, as required.

- Failure to complete the mandatory annual training is grounds for suspension or termination of all physical or electronic access privileges and removal from work on the contract until such time as the required training is complete.

SECURITY INCIDENT INVESTIGATION

- The Contractor, Subcontractor, their employees, or business associates shall immediately (within one hour) report suspected security / privacy incidents to the VA OIT’s Enterprise Service Desk (ESD) by calling (855) 673-4357 (TTY: 711). The ESD is OIT’s 24/7/365 single point of contact for IT-related issues. After reporting to the ESD, the Contractor, Subcontractor, their employees, or business associates shall, within one hour, provide the COR/CO the incident number received from the ESD.

- To the extent known by the Contractor/Subcontractor, the Contractor/Subcontractor's notice to VA shall identify the information involved and the circumstances surrounding the incident, including the following:

  - The date and time (or approximation of) the Security Incident occurred.

  - The names of individuals involved (when applicable).

  - The physical and logical (if applicable) location of the incident.

  - Why the Security Incident took place (i.e., catalyst for the failure).

  - The amount of data belonging to VA believed to have been compromised.

  - The remediation measures the Contractor is taking to ensure no future incidents of a similar nature.

- After the Contractor has provided the initial detailed incident summary to VA, they will continue to provide written updates on any new and relevant circumstances or facts they discover. The Contractor, Subcontractor, and their employes shall fully cooperate with VA or third-party entity performing an independent risk analysis on behalf of VA. Failure to cooperate may be deemed a material breach and grounds for contract termination.

- VA IT Contractors shall follow VA Handbook 6500, Risk Management Framework for VA Information Systems VA Information Security Program, and VA Information Security Knowledge Service guidance for implementing an Incident Response Plan or integrating with an existing VA implementation.

- In instances of theft or break-in or other criminal activity, the Contractor/Subcontractor must concurrently report the incident to the appropriate law enforcement entity (or entities) of jurisdiction, including the VA OIG, and the VA Office of Security and Law Enforcement. The Contractor, its employees, and its Subcontractors and their employees shall cooperate with VA and any law enforcement authority responsible for the investigation and prosecution of any possible criminal law violation(s) associated with any incident. The Contractor/Subcontractor shall cooperate with VA in any civil litigation to recover VA information, obtain monetary or other compensation from a third party for damages arising from any incident, or obtain injunctive relief against any third party arising from, or related to, the incident.

- The Contractor shall comply with VA Handbook 6500.2, Management of Breaches Involving Sensitive Personal Information, which establishes the breach management policies and assigns responsibilities for the oversight, management and reporting procedures associated with managing of breaches.

- With respect to unsecured Protected Health Information (PHI), the Contractor is deemed to have discovered a data breach when the Contractor knew or should have known of breach of such information. When a business associate is part of VHA contract, notification to the covered entity (VHA) shall be made in accordance with the executed BAA.

- If the Contractor or any of its agents fails to protect VA sensitive personal information or otherwise engages in conduct which results in a data breach involving any VA sensitive personal information the Contractor/Subcontractor processes or maintains under the contract; the Contractor shall pay liquidated damages to the VA as set forth in clause 852.211-76, Liquidated Damages—Reimbursement for Data Breach Costs.

SPECIALIZED DEVICES/SYSTEMS (MEDICAL DEVICES, SPECIAL PURPOSE SYSTEMS, RESEARCH SCIENTIFIC COMPUTING)

- Contractor supplies/delivered Medical Devices, Special Purpose Systems-Operational Technology (SPS-OT) and Research Scientific Computing Devices shall comply with all applicable Federal law, regulations, and VA policies. New developments require creation, testing, evaluation, and authorization in compliance with processes specified on the Specialized Device Cybersecurity Department Enterprise Risk Management (SDCD-ERM) Portal, VA Directive 6550, Pre-Procurement Assessment and Implementation of Medical Devices/Systems, VA Handbook 6500, and the VA Information Security Knowledge Service. Deviations from Federal law, regulations, and VA Policy are identified and documented as part of VA Directive 6550 and/or the VA Enterprise Risk Analysis (ERA) processes for Specialized Devices/Systems processes.

- All Contractors and third-party service providers shall address and/or integrate applicable VA Handbook 6500 and Information Security Knowledge Service specifications in delivered IT systems/solutions, products and/or services. If systems/solutions, products and/or services do not directly match VA security requirements, the Contractor shall work though the COR/CO for governance or resolution.

- The Contractor shall certify to the COR/CO that devices/systems that have completed the VA Enterprise Risk Analysis (ERA) process for Specialized Devices/Systems are fully functional and operate correctly as intended. Devices/systems must follow the VA ERA authorized configuration prior to acquisition and connection to the VA computing environment. If VA determines a new VA ERA needs to be created, the Contractor shall provide required technical support to develop the configuration settings. Major changes to a previously approved device/system will require a new ERA.

- The Contractor shall comply with all practices documented by the Food Drug and Administration (FDA) Premarket Submission for Management of Cybersecurity in Medical Devices and Postmarket Management of Cybersecurity in Medical Devices.

- The Contractor shall design devices capable of accepting all applicable security patches with or without the support of the Contractor personnel. If patching can only be completed by the Contractor, the Contractor shall commit the resources needed to patch all applicable devices at all VA locations. If unique patching instructions or packaging is needed, the Contractor shall provide the necessary information in conjunction with the validation/testing of the patch. The Contractor shall apply security patches within 30 business days of the patch release and have a formal tracking process for any security patches not implemented to include explanation when a device cannot be patched.

- The Contractor shall provide devices able to install and maintain VA-approved antivirus capabilities with the capability to quarantine files and be updated as needed in response to incidents. Alternatively, a VA-approved whitelisting application may be used when the Contractor cannot install an anti-virus / anti-malware application.

- The Contractor shall verify and document all SOFTWARE embedded within the device does not contain any known viruses or malware before delivery to or installation at a VA location.

- Devices and other equipment or systems containing media (hard drives, optical disks, solid state, and storage via chips/firmware) with VA sensitive information will be returned to the Contractor with media removed. When the contract requires return of equipment, the options available to the Contractor are the following:

  - The Contractor shall accept the system without the drive, firmware and solid state.

  - VA’s initial device purchase includes a spare drive or other replacement media which must be installed in place of the original drive at time of turn-in; or

  - Due to the highly specialized and sometimes proprietary hardware and SOFTWARE associated with the device, if it is not possible for VA to retain the hard drive, firmware, and solid state, then:

    - The equipment Contractor shall have an existing BAA if the device being traded in has sensitive information stored on it and hard drive(s) from the system are being returned physically intact.

    - Any fixed hard drive, Complementary Metal-Oxide-Semiconductor (CMOS), Programmable Read-Only Memory (PROM), solid state and firmware on the device must be non-destructively sanitized to the greatest extent possible without negatively impacting system operation. Selective clearing down to patient data folder level is recommended using VA approved and validated overwriting technologies/methods/tools. Applicable media sanitization specifications need to be pre-approved and described in the solicitation, contract, or order.