DESCRIPTION

- Title: Automated Immunology Analyzer System at Central Arkansas Veterans Healthcare System (CAVHS), AR, 72205. Brand Name or Equal.

- Purpose:

- The Network Contracting Office (NCO) 16 located at 5075 Westheimer Road,

Galleria Financial Center, Suite 750, Houston, Texas, hereby issues the following Sources Sought/ Request for Information (RFI) with the intent of determining the contractor interest, capabilities, qualifications of potential businesses, nonmanufacturer rule compliance, Buy American Act compliance, Trade Agreement compliance and estimated industry pricing information for market research and estimated budgetary purpose for Contractor to assist with the BioPlex 2200 Analyzer System, Reagents, Controls, Calibrators, Consumables, Installation, Validation, Training, Maintenance, and Technical Support for Central Arkansas Veterans Healthcare System (CAVHS), AR, 72205.

- Notice of Intent to Sole Source.  If there are no responses to this notice, a determination by the Government not to compete based upon responses to this notice is solely within the discretion of the Government to sole source the requirement. No response is a response.

- Government request that interested offerors complete all questions below, provide required documentation, descriptive literature and authorization letter as described below. The Government is in no way obligated to do business with or to enter into any form of contract with any person, firm or other entity that receives or responds to this announcement.

- Objective: To find qualified and certified contractors with the capability to provide the following:

- The Contractor shall provide all equipment, reagents, controls, calibrators, consumable parts, software, labor, installation, validation, training, and ongoing maintenance and technical support necessary to furnish fully functional BioPlex 2200 analyzer system at CAVHS. Brand Name or Equal.

- Delivery Location:

Primary place of performance:

Central Arkansas Veterans Healthcare System (CAVHS)

Pathology and Laboratory Medicine Service

4300 West 7th Street, Room 2E 155 LR

Little Rock, Arkansas 72205

- Period of Performance:

Five (5) Year IDIQ

- Responses Requested: The CAVHS requires all the following information and questions answered in this RFI/SS: Questions and information that is not provided shall be considered nonresponsive to the Request for Information and contractor shall not be considered as part of the market research.

- Contractors shall provide their point(s) of contact name, address, telephone number, and email address, and the company's business size, and SAM Unique Entity Identification Number.

- Is your company a small business, SDBs, HUBZone, or 8A concern? Please provide proof of qualifications.

- Does your company meet all required four prong factors for the Nonmanufacturer Rule (NMR) per guidance of SBA and 13 C.F.R. § 121.406? Offer must answer all questions below.

- Does your company exceed the 500-employee alternative size standard for nonmanufacturers?

- Is your company primarily engaged in the retail or wholesale trade and normally sells the type of products being supplied?

- Does your company must take ownership or possession of the item(s) with its personnel, equipment, or facilities in a manner consistent with industry practice?

- Does your company supply the end products of a small business manufacturer or processor made in the United States?

- Is your company considered small under the NAICS code 334516.

- Do you (offer) manufacture the product yourself or a distributor of the items being researched above?

- Are you an authorized distributor/reseller? Provide evidence from original equipment manufacturer (OEM).

- Do you (offer) supply a product manufactured by another domestic small business, if it is a nonmanufacturer?

- Do you (offer) supply the product of any sized manufacturer if SBA has granted a waiver to the nonmanufacturer rule?

- Is your (offer) Buy American Certificate completed in SAM? Please provide evidence with response.

- Is your (offer) Trade Agreement Certificate completed in SAM?  Please provide evidence with response.

- Information for the name brand or equal product being requested, is it manufactured outside the US-Domestic non-available?

- Information for the name brand or equal product being requested, is it a foreign manufactured end product under the trade agreement?

- What is your schedule of delivery after receipt of order for the name brand or equal product?

- For the name brand or equal product required, Is your company available under any;

Government Wide Agency Contract (GWAC)

General Services Administration Schedules (GSA)

Indefinite Delivery Indefinite Quantity (IDIQ)

Blanket Purchase Agreement (BPA)

If so, please list the contract number for name brand or equal product.

- Provide descriptive literature demonstrating ability to meet all of the requirement/salient characteristics listed if other than brand name product is being offered. Please provide list of contracts for commercial, federal, state, and local governments. Listed contracts shall provide communication on the government requirement illustrating the capability for comparison. to the name brand or equal product meeting the specific requirements per the draft Statement of Work.

- Contractors may provide an estimated industry pricing for name brand or equal product for market research and budgetary purpose.

- Opportunity:

- NCO 16, is seeking information from potential contractors on their ability to

provide this requirement. THIS IS A NOTICE OF INTENT TO SOLE SOURCE and SOURCES SOUGHT FOR REQUEST INFORMATION (RFI) ONLY. Small Business Concerns are encouraged to provide responses to this Sources Sought in order to assist the CAVHS in determining potential levels of competition available in the industry. Contractor shall possess the capability to provide all requirements and objectives.

- Instructions and Response Guidelines:

- Sources Sought responses are due by 03/07/2026 at 3:00 PM (CST) via email to t.miller@va.gov. Telephone requests or inquires will not be accepted. Public Information will not be provided and shall be utilized to the full extent possible. This a New Requirement.

- The subject line shall read: 36C25626Q0584 - Sources Sought – Automated Immunology Analyzer System.  NO SOLICITATION EXISTS AT THIS TIME. There is no page limitation on subparagraphs 6(a) - 6(c17).

- Please provide the information you deem relevant in order to respond to the specific inquiries of the Sources Sought. Information provided will be used solely by CAVHS as "market research" and will not be released outside of the CAVHS Purchasing and Contract Team.  This Sources Sought notice does not constitute a Request for Proposal (RFP), Invitation for Bid (IFB), or Request for Quotation (RFQ), and it is not to be construed as a commitment by the Government to enter into a contract, nor will the Government pay for the information submitted in response to this request. All information contained in this Sources Sought is preliminary as well as subject to modification and is in no way binding on the Government.

- In accordance with FAR 15.201(e), responses to this notice are not offers and cannot be accepted by the U.S. Government to form a binding contract. If a solicitation is released, it will be synopsized on Contracting Opportunities (https://beta.sam.gov/) website or GSA. It is the responsibility of the interested parties to monitor these sites for additional information pertaining to this Sources Sought, or future RFP/RFQ.

- Contact Information:

Contracting Specialist, T Miller

Email address: t.miller@va.gov

- THIS NOTICE IS NOT A REQUEST FOR COMPETITIVE QUOTES/PROPOSALS; however, any firm that believes it can meet the requirements may provide answers, responses and give written notification prior to the response due date and time. Supporting evidence and responses.

Your responses to this notice are appreciated.

See Statement of Work (SOW) Below:

STATEMENT OF WORK

Bio-Rad BioPlex 2200 Brand Name or Equal

Central Arkansas Veterans Healthcare System (CAVHS)

Little Rock, Arkansas

- Contract Title:

- BioPlex 2200 Analyzer System, Reagents, Controls, Calibrators, Consumables, Installation, Validation, Training, Maintenance, and Technical Support – Indefinite Delivery, Indefinite Quantity (IDIQ) Contract, Base Period plus Four (4) Option Years.

- Background:

- The Central Arkansas Veterans Healthcare System (CAVHS), Pathology and Laboratory Medicine Service (PLMS), requires FDA-cleared BioPlex 2200 analyzer and associated reagents, controls, calibrators, and consumable parts to support high-complexity laboratory testing and to improve operational efficiency and patient outcomes. The BioPlex 2200 systems will integrate with the Veterans Health Information Systems and Technology Architecture (VistA/CPRS) and/or the enterprise electronic health record to support automated result transmission and standardized reporting workflows.

- Current testing for the affected assays is performed using a combination of manual and semi-automated platforms, which limits throughput, increases hands-on time, and can delay result turnaround times. Implementing BioPlex 2200 analyzer and associated test menus under an IDIQ structure will allow CAVHS to address fluctuating test volumes while ensuring continuity of operations, quality, and regulatory compliance across the period of performance.

- Scope of Work:

- The Contractor shall provide all equipment, reagents, controls, calibrators, consumable parts, software, labor, installation, validation, training, and ongoing maintenance and technical support necessary to furnish fully functional BioPlex 2200 analyzer system at CAVHS and at any additional VA facilities authorized to order under this IDIQ.

- BioPlex 2200 Systems, Reagents, and Consumables:

    - Provide FDA-cleared BioPlex 2200 analyzer sized and configured for the test menus and workloads specified in each task order.

    - Provide all associated reagents, assay kits, controls, calibrators, and consumables (including but not limited to cuvettes, tubing, probe wash, reaction vessels, instrument-specific filters, waste containers, and other consumable parts) required to perform testing over the term of the contract.

    - Ensure lot-to-lot continuity and ship new lots in time to allow parallel testing during lot-to-lot verification; provide documentation for each lot.

    - Provide current Instructions for Use, Safety Data Sheets (SDS), and package inserts for all products in print and/or electronic format.

- System Installation and Integration:

    - Deliver, uncrate, install, and configure the BioPlex 2200 analyzer and associated hardware at the designated laboratory locations within CAVHS.

    - Configure instruments to match local test codes, reporting units, critical values, reference intervals, and reflex rules as specified by PLMS and the Contracting Officer’s Representative (COR).

    - Implement and test interfaces between the BioPlex 2200 analyzer and the laboratory information system (LIS) and/or middleware, including mapping of orders, results, comments, flags, and QC data.

    - Verify that all installations comply with manufacturer requirements, electrical and environmental conditions, and facility safety, infection control, and occupational health requirements.

- Method Validation and Verification:

    - Perform or support installation qualification (IQ), operational qualification (OQ), and performance qualification (PQ) as applicable to each instrument and assay.

    - Assist PLMS with method validation/verification studies in accordance with CLIA, CAP, and VA standards, including accuracy, precision, linearity, reportable range, reference interval verification, and method comparison where required.

    - Provide necessary validation materials, guidance documents, and recommended protocols; generate summary reports suitable for inclusion in laboratory validation files and accreditation inspections.

- System Maintenance and Technical Support:

    - Provide full-service maintenance for all covered BioPlex 2200 analyzer and associated modules, including preventive maintenance, corrective maintenance, service parts, and labor.

    - Provide telephone technical support 24 hours per day, 7 days per week, 365 days per year for instrument and assay issues.

    - Provide remote diagnostic support capabilities when permitted by VA information security policies and approved by the COR.

    - Perform preventive maintenance at manufacturer-recommended intervals, coordinated with PLMS to minimize disruption to patient testing.

    - Provide corrective maintenance with response times that meet or exceed the performance standards in Section 7.

  - Professional Services (As Needed)

    - Provide on-site technical support for complex troubleshooting, workflow optimization, and implementation of new assays or software upgrades.

    - Assist with development and refinement of instrument rules, error-handling parameters, QC rules, and reflex testing algorithms in collaboration with PLMS and LIS staff.

    - Support validation of new tests, parameter changes, and software or firmware updates throughout the contract period.

  - Training and Competency Support

    - Provide initial and ongoing operator training for CAVHS laboratory staff on BioPlex 2200 system operation, routine and advanced maintenance, troubleshooting, QC management, and reagent handling.

    - Provide training materials, competency checklists, and documentation suitable for use in competency assessments and accreditation audits.

    - Provide refresher training and training for new staff when requested by the COR; training may be on-site and/or virtual, as agreed by the COR and Contractor.

- Specific Tasks and Deliverables:

  - Customer Service Requirements

    - The Contractor shall:

    - Provide 24/7/365 telephone technical support for urgent instrument and assay issues.

    - Designate primary and backup technical contacts for escalation by the COR and Laboratory Information Manager.

    - Respond to critical issues (e.g., instrument down, inability to perform patient testing) within 2–4 hours of notification and to non-critical issues no later than the next business day.

    - Document all service requests, actions taken, and final resolution, and provide copies or access to service records upon request by the COR.

    - Participate, upon request, in periodic performance and customer service review meetings with PLMS.

  - Product Support and Changes

    - Immediately notify the COR of any product recalls, product notices, safety alerts, or mandatory software/firmware updates related to BioPlex 2200 analyzer, reagents, controls, or consumables.

    - Not substitute analyzer, reagents, controls, calibrators, or consumables without prior written approval from the COR and the Contracting Officer.

    - Ensure all changes to contract terms, conditions, or pricing are formally documented and signed by the VA Contracting Officer.

    - Replace and/or reimburse recalled or defective products at no cost to the Government, including providing replacement materials in time to avoid interruption of patient testing where practicable.

  - Site Conditions and Access

    - Perform delivery, installation, on-site support, and training during normal clinical business hours (Monday–Friday, 8:00 AM – 4:30 PM Central Time, excluding federal holidays and weekends), unless otherwise authorized by the COR.

    - Ensure all Contractor staff comply with facility security, safety, infection control, and dress code requirements, including restrictions on food, drink, and cell phone use in the laboratory.

    - Comply with CAVHS’s smoke-free campus policy.

  - Equipment Protection and Property

    - Protect all Government property, facilities, and equipment from damage during all phases of work.

    - Repair or replace, at Contractor expense, any Government property damaged as a result of Contractor or subcontractor actions.

  - System Security Requirements

    - Comply with all applicable VA records management, privacy, and information security laws, regulations, and policies, including but not limited to the Federal Records Act, Privacy Act, Freedom of Information Act, and applicable VA/VHA directives and handbooks.

    - Access Protected Health Information (PHI) only as necessary to perform contract services and only under secure, approved conditions.

    - Treat all data created for Government use in connection with instrument connectivity, testing, or reporting as Federal records under Government ownership and control; do not remove or destroy such records except as authorized under VA records schedules.

    - Not remove hard drives, non-volatile memory units, or other data-bearing components from VA facilities without written approval from the COR and Contracting Officer.

    - Execute a Business Associate Agreement (BAA) if Contractor personnel will access PHI, and comply with all applicable VA/VHA privacy requirements.

  - Compliance and Training Requirements

    - Ensure Contractor employees who create, handle, or access Federal records or VA sensitive information complete required VA-provided records management, privacy, and cybersecurity training, including periodic refresher training.

    - Comply with relevant VA and VHA directives and handbooks governing PLMS operations, privacy, cybersecurity, and records management (e.g., VA Directives 1106, 6500, 6502; VHA Directive 1605), as updated during the contract period.

  - Delivery and Installation

    - Deliver and install BioPlex 2200 analyzer and related equipment at:

Central Arkansas Veterans Healthcare System (CAVHS)
Pathology and Laboratory Medicine Service
4300 West 7th Street, Room 2E-155-LR
Little Rock, Arkansas 72205

    - and at any additional locations identified in individual task orders.

    - Coordinate delivery and installation timelines with the COR; conduct a pre-installation meeting as needed to verify site readiness, utilities, space, and interface requirements.

    - Conduct a joint inspection with the COR upon installation to confirm that all equipment, accessories, and documentation are delivered and functional.

    - Provide manufacturer-recommended maintenance schedules, operator manuals, and other operational documentation in binder(s) and/or electronic format.

- PLACE OF PERFORMANCE

Primary place of performance:

Central Arkansas Veterans Healthcare System (CAVHS)
Pathology and Laboratory Medicine Service
4300 West 7th Street, Room 2E-155-LR
Little Rock, Arkansas 72205

  - Additional VA facilities may place orders under this IDIQ as authorized; each task order will identify the specific place(s) of performance.

- PERIOD OF PERFORMANCE

  - Base Period: August 1, 2026 through July 11, 2027.

  - Option Years: Four (4) option years of one (1) year each, exercisable at the sole discretion of the Government and subject to the availability of funds.

  - The overall ordering period shall not exceed the combined duration of the base period plus four (4) option years. An annual review of services and performance shall be conducted by the COR and Contracting Office at least 90 days prior to the anniversary date or option exercise.

- PERFORMANCE MONITORING AND ACCEPTANCE

  - Acceptance Procedures

    - The COR shall verify that delivered analyzer, reagents, consumables, and services meet all specifications described in this SOW and in applicable task orders prior to acceptance.

    - Instruments may not be placed into clinical use until required validation/verification studies are completed and approved by PLMS.

    - Any disputes regarding work quality or performance shall be referred to the Contracting Officer for final resolution.

  - Performance Monitoring Criteria

    - Vendor performance shall be monitored based on:

    - Instrument uptime and reliability (target ≥ 95–98% uptime per quarter, excluding scheduled preventive maintenance).

    - Timeliness and effectiveness of technical support and on-site service responses.

    - Ability to maintain uninterrupted supply of reagents, controls, calibrators, and consumables without stock-outs that affect patient care.

    - Quality and timeliness of training and professional services.

    - Compliance with information security, privacy, and records management requirements.

  - Quality Assurance Metrics

    - Initial response to critical support inquiries within 2–4 hours, with at least 95% compliance.

    - Preventive maintenance completed within agreed timeframes and documented appropriately.

    - Annual or more frequent performance review, including Contractor Performance Assessment Reporting System (CPARS) evaluations as applicable; exercise of option years may be contingent upon satisfactory performance.

- GOVERNMENT FURNISHED EQUIPMENT (GFE) / INFORMATION (GFI)

  - The Government will furnish space, utilities, and network connectivity as needed for installation and operation of BioPlex 2200 systems at CAVHS and other ordering facilities. No other GFE or GFI is anticipated; the Contractor is responsible for all tools, service parts, test materials, and other resources required to perform this contract.

- OTHER PERTINENT INFORMATION

  - The Contractor may require proprietary access codes and service tools to perform maintenance on BioPlex 2200 analyzer; such access shall be managed in accordance with VA information security requirements.

  - Contractor and subcontractor personnel are subject to VA background investigation and badging requirements appropriate to the level of system and data access required.

  - All employees, contractors, and visitors shall comply with local infection control procedures per VA and CAVHS policy.

- FEDERAL HOLIDAYS

  - Work will not ordinarily be performed on the following Federal holidays, except when specifically authorized by the COR for urgent service: New Year’s Day, Martin Luther King Jr.’s Birthday, Presidents Day, Memorial Day, Juneteenth, Independence Day, Labor Day, Columbus Day, Veterans Day, Thanksgiving Day, and Christmas Day. When a holiday falls on Sunday, the following Monday is observed; when a holiday falls on Saturday, the preceding Friday is observed.

- PRICE/COST AND ORDERING STRUCTURE (IDIQ)

  - This contract shall be structured as Reagent Rental and Indefinite Delivery, Indefinite Quantity (IDIQ) contract. Individual task orders will define the specific BioPlex 2200 models/configurations, estimated annual testing volumes, and required quantities of reagents, controls, calibrators, and consumables, as well as any associated installation, validation, training, and maintenance services.

  - e Contractor’s pricing schedule shall, at a minimum, include:

    - Analyzer placement and, if applicable, cost-per-test or reagent-rental options.

    - Unit prices for reagents, controls, calibrators, and key consumables.

    - Service and maintenance coverage, if separately priced.

    - Rates for professional services, implementation, and training support.

- CONTRACTING OFFICER AND CONTRACTING OFFICER’S REPRESENTATIVE

  - Contracting Officer (CO): To be designated by the VA Contracting Activity.

  - Contracting Officer’s Representative (COR): To be designated by CAVHS; contact information will be identified in the contract award and/or task orders.

  - All official contract communications shall be directed through the CO and COR, as appropriate.

14. RECORDS MANAGEMENT OBLIGATIONS

A.  Applicability

This clause applies to all Contractors whose employees create, work with, or otherwise handle Federal records, as defined in Section B, regardless of the medium in which the record exists.

B.  Definitions

“Federal record” as defined in 44 U.S.C. § 3301, includes all recorded information, regardless of form or characteristics, made or received by a Federal agency under Federal law or in connection with the transaction of public business and preserved or appropriate for preservation by that agency or its legitimate successor as evidence of the organization, functions, policies, decisions, procedures, operations, or other activities of the United States Government or because of the informational value of data in them.

The term Federal record:

- includes [Agency] records.

- does not include personal materials.

- applies to records created, received, or maintained by Contractors pursuant to their [Agency] contract.

- may include deliverables and documentation associated with deliverables.

C.  Requirements

- Contractor shall comply with all applicable records management laws and regulations, as well as National Archives and Records Administration (NARA) records policies, including but not limited to the Federal Records Act (44 U.S.C. chs. 21, 29, 31, 33), NARA regulations at 36 CFR Chapter XII Subchapter B, and those policies associated with the safeguarding of records covered by the Privacy Act of 1974 (5 U.S.C. 552a). These policies include the preservation of all records, regardless of form or characteristics, mode of transmission, or state of completion.

- In accordance with 36 CFR 1222.32, all data created for Government use and delivered to, or falling under the legal control of, the Government are Federal records subject to the provisions of 44 U.S.C. chapters 21, 29, 31, and 33, the Freedom of Information Act (FOIA) (5 U.S.C. 552), as amended, and the Privacy Act of 1974 (5 U.S.C. 552a), as amended and must be managed and scheduled for disposition only as permitted by statute or regulation.

- In accordance with 36 CFR 1222.32, Contractor shall maintain all records created for Government use or created in the course of performing the contract and/or delivered to, or under the legal control of the Government and must be managed in accordance with Federal law. Electronic records and associated metadata must be accompanied by sufficient technical documentation to permit understanding and use of the records and data.

- [Agency] and its contractors are responsible for preventing the alienation or unauthorized destruction of records, including all forms of mutilation. Records may not be removed from the legal custody of [Agency] or destroyed except for in accordance with the provisions of the agency records schedules and with the written concurrence of the Head of the Contracting Activity. Willful and unlawful destruction, damage or alienation of Federal records is subject to the fines and penalties imposed by 18 U.S.C. 2701. In the event of  any unlawful or accidental removal, defacing, alteration, or destruction of records, Contractor must report to [Agency]. The agency must report promptly to NARA in accordance with 36 CFR 1230.

- The Contractor shall immediately notify the appropriate Contracting Officer upon discovery of any inadvertent or unauthorized disclosures of information, data, documentary materials, records or equipment. Disclosure of non-public information is limited to authorized personnel with a need-to-know as described in the [contract vehicle]. The Contractor shall ensure that the appropriate personnel, administrative, technical, and physical safeguards are established to ensure the security and confidentiality of this information, data, documentary material, records and/or equipment is properly protected. The Contractor shall not remove material from Government facilities or systems, or facilities or systems operated or maintained on the Government’s behalf, without the express written permission of the Head of the Contracting Activity. When information, data, documentary material, records and/or equipment is no longer required, it shall be returned to [Agency] control or the Contractor must hold it until otherwise directed. Items returned to the Government shall be hand carried, mailed, emailed, or securely electronically transmitted to the Contracting Officer or address prescribed in the [contract vehicle]. Destruction of records is EXPRESSLY PROHIBITED unless in accordance with Paragraph (4).

- The Contractor is required to obtain the Contracting Officer's approval prior to engaging in any contractual relationship (sub-contractor) in support of this contract requiring the disclosure of information, documentary material and/or records generated under, or relating to, contracts. The Contractor (and any sub-contractor) is required to abide by Government and [Agency] guidance for protecting sensitive, proprietary information, classified, and controlled unclassified information.

- The Contractor shall only use Government IT equipment for purposes specifically tied to or authorized by the contract and in accordance with [Agency] policy.

- The Contractor shall not create or maintain any records containing any non-public [Agency] information that are not specifically tied to or authorized by the contract.

- The Contractor shall not retain, use, sell, or disseminate copies of any deliverable that contains information covered by the Privacy Act of 1974 or that which is generally protected from public disclosure by an exemption to the Freedom of Information Act.

- The [Agency] owns the rights to all data and records produced as part of this contract. All deliverables under the contract are the property of the U.S. Government for which [Agency] shall have unlimited rights to use, dispose of, or disclose such data contained therein as it determines to be in the public interest. Any Contractor rights in the data or deliverables must be identified as required by FAR 52.227-11 through FAR 52.227-20.

- Training.  All Contractor employees assigned to this contract who create, work with, or otherwise handle records are required to take [Agency]-provided records management training. The Contractor is responsible for confirming training has been completed according to agency policies, including initial training and any annual or refresher training.

[Note: To the extent an agency requires contractors to complete records management training, the agency must provide the training to the contractor.]

D.  Flow down of requirements to subcontractors

- The Contractor shall incorporate the substance of this clause, its terms and requirements including this paragraph, in all subcontracts under this [contract vehicle], and require written subcontractor acknowledgment of same.

- Violation by a subcontractor of any provision set forth in this clause will be attributed to the Contractor.

15. INFORMATION SECURITY

- GENERAL. This entire section applies to all acquisitions requiring any Information Security and Privacy language. Contractors, contractor personnel, subcontractors and subcontractor personnel will be subject to the same federal laws, regulations, standards, VA directives and handbooks, as VA personnel regarding information and information system security and privacy.

- VA INFORMATION CUSTODIAL LANGUAGE. This entire section applies to all acquisitions requiring any Information Security and Privacy language.

  - The Government shall receive unlimited rights to data/intellectual property first produced and delivered in the performance of this contract or order (hereinafter “contract”) unless expressly stated otherwise in this contract. This includes all rights to source code and all documentation created in support thereof. The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General. The primary clause used to define computer software license (not data/intellectual property first produced under this contractor or order) is FAR 52.227-19, Commercial Computer Software License.

  - Information made available to the contractor by VA for the performance or administration of this contract will be used only for the purposes specified in the service agreement, SOW, PWS, PD, and/or contract. The contractor shall not use VA information in any other manner without prior written approval from a VA Contracting Officer (CO). The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General.

  - VA information will not be co-mingled with any other data on the contractor’s information systems or media storage systems. The contractor shall ensure compliance with Federal and VA requirements related to data protection, data encryption, physical data segregation, logical data segregation, classification requirements and media sanitization.

  - VA reserves the right to conduct scheduled or unscheduled audits, assessments, or investigations of contractor Information Technology (IT) resources to ensure information security is compliant with Federal and VA requirements. The contractor shall provide all necessary access to records (including electronic and documentary materials related to the contracts and subcontracts) and support (including access to contractor and subcontractor staff associated with the contract) to VA, VA's Office Inspector General (OIG),

and/or Government Accountability Office (GAO) staff during periodic control assessments, audits, or investigations.

  - The contractor may only use VA information within the terms of the contract and applicable Federal law, regulations, and VA policies. If new Federal information security laws, regulations or VA policies become applicable after execution of the contract, the parties agree to negotiate contract modification and adjustment necessary to implement the new laws, regulations, and/or policies.

  - The contractor shall not make copies of VA information except as specifically authorized and necessary to perform the terms of the contract. If copies are made for restoration purposes, after the restoration is complete, the copies shall be destroyed in accordance with VA Directive 6500, VA Cybersecurity Program and VA Information Security Knowledge Service.

  - If a Veterans Health Administration (VHA) contract is terminated for default or cause with a business associate, the related local Business Associate Agreement (BAA) shall also be terminated and actions taken in accordance with VHA Directive 1605.05, Business Associate Agreements. If there is an executed national BAA associated with the contract, VA will determine what actions are appropriate and notify the contactor.

  - The contractor shall store and transmit VA sensitive information in an encrypted form, using VA-approved encryption tools which are, at a minimum, Federal Information Processing Standards (FIPS) 140-2, Security Requirements for Cryptographic Modules (or its successor) validated and in conformance with VA Information Security Knowledge Service requirements. The contractor shall transmit VA sensitive information using VA approved Transport Layer Security (TLS) configured with FIPS based cipher suites in conformance with National Institute of Standards and Technology (NIST) 800-52, Guidelines for the Selection, Configuration and Use of Transport Layer Security (TLS) Implementations.

  - The contractor’s firewall and web services security controls, as applicable, shall meet or exceed VA’s minimum requirements.

  - Except for uses and disclosures of VA information authorized by this contract for performance of the contract, the contractor may use and disclose VA information only in two situations: (i) in response to a qualifying order of a court of competent jurisdiction after notification to VA CO (ii) with written approval from the VA CO. The contractor shall refer all requests for, demands for production of or inquiries about, VA information and information systems to the VA CO for response.

  - Notwithstanding the provision above, the contractor shall not release VA records protected by Title 38 U.S.C. § 5705, Confidentiality of medical quality-assurance records and/or Title 38 U.S.C. § 7332, Confidentiality of certain medical records pertaining to drug addiction, sickle cell anemia, alcoholism or alcohol abuse or infection with Human Immunodeficiency Virus (HIV). If the contractor is in receipt of a court order or other requests for the above-mentioned information, the contractor shall immediately refer such court order or other requests to the VA CO for response.

  - Information made available to the contractor by VA for the performance or administration of this contract or information developed by the contractor in performance or administration of the contract will be protected and secured in accordance with VA Directive 6500 and Identity and Access Management (IAM) Security processes specified in the VA Information Security Knowledge Service.

  - Any data destruction done on behalf of VA by a contractor shall be done in accordance with National Archives and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records and Information Management, VA Handbook 6300.1, Records Management Procedures, and applicable VA Records Control Schedules.

  - The contractor shall provide its plan for destruction of all VA data in its possession according to VA Directive 6500 and NIST 800-88, Guidelines for Media Sanitization prior to termination or completion of this contract. If directed by the COR/CO, the contractor shall return all Federal Records to VA for disposition.

  - Any media, such as paper, magnetic tape, magnetic disks, solid state devices or optical discs that is used to store, process, or access VA information that cannot be destroyed shall be returned to VA.The contractor shall hold the appropriate material until otherwise directed by the Contracting Officer’s Representative (COR) or CO. Items shall be returned securely via VA-approved methods. VA sensitive information must be transmitted utilizing VA-approved encryption tools which are validated under FIPS 140-2 (or its successor) and NIST 800-52. If mailed, the contractor shall send via a trackable method (USPS, UPS, FedEx, etc.) and immediately provide the COR/CO with the tracking information. Self-certification by the contractor that the data destruction requirements above have been met shall be sent to the COR/CO within 30 business days of termination of the contract.

  - All electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) used to store, process or access VA information will not be returned to the contractor at the end of lease, loan, or trade-in. Exceptions to this paragraph will only be granted with the written approval of the VA CO.

- ACCESS TO VA INFORMATION AND VA INFORMATION SYSTEMS. This section

applies when any person requires access to information made available to the contractor by VA for the performance or administration of this contract or information developed by the contractor in performance or administration of the contract.

  - A contractor/subcontractor shall request logical (technical) or physical access to VA information and VA information systems for their employees and subcontractors only to the extent necessary to perform the services specified in the solicitation or contract. This includes indirect entities, both affiliate of contractor/subcontractor and agent of contractor/subcontractor.

  - Contractors and subcontractors shall sign the VA Information Security Rule of Behavior (ROB) before access is provided to VA information and information systems (see Section 4, Training, below). The ROB contains the minimum user compliance requirements and does not supersede any policies of VA facilities or other agency components which provide higher levels of protection to VA’s information or information systems. Users who require privileged access shall complete the VA elevated privilege access request processes before privileged access is granted.

  - All contractors and subcontractors working with VA information are subject to the same security investigative and clearance requirements as those of VA appointees or employees who have access to the same types of information. The level and process of background security investigations for contractors shall be in accordance with VA Directive and Handbook 0710, Personnel Suitability and Security Program. The Office of Human Resources and Administration/Operations, Security and Preparedness (HRA/OSP) is responsible for these policies and procedures. Contract personnel who require access to classified information or information systems shall have an appropriate security clearance. Verification of a Security Clearance shall be processed through the Special Security Officer located in HRA/OSP. Contractors shall conform to all requirements stated in the National Industrial Security Program Operating Manual (NISPOM).

  - All contractors and subcontractors shall comply with conditions specified in VAAR 852.204-71(d); Contractor operations required to be in United States. All contractors and subcontractors working with VA information must be permanently located within a jurisdiction subject to the law of the United States or its Territories to the maximum extent feasible. If services are proposed to be performed abroad the contractor must state where all non-U.S. services are provided. The contractor shall deliver to VA a detailed plan specifically addressing communications, personnel control, data protection and potential legal issues. The plan shall be approved by the COR/CO in writing prior to access being granted.

  - The contractor shall notify the COR/CO in writing immediately (no later than 24 hours) after personnel separation or occurrence of other causes. Causes may include the following:

    - Contractor/subcontractor personnel no longer has a need for access to VA information or VA information systems.

    - Contractor/subcontractor personnel are terminated, suspended, or otherwise has their work on a VA project discontinued for any reason.

    - Contractor believes their own personnel or subcontractor personnel may pose a threat to their company’s working environment or to any company-owned property. This includes contractor-owned assets, buildings, confidential data, customers, employees, networks, systems, trade secrets and/or VA data.

    - Any previously undisclosed changes to contractor/subcontractor background history are brought to light, including but not limited to changes to background investigation or employee record.

    - Contractor/subcontractor personnel have their authorization to work in the United States revoked.

    - Agreement by which contractor provides products and services to VA has either been fulfilled or terminated, such that VA can cut off electronic and/or physical access for contractor personnel.

  - In such cases of contract fulfillment, termination, or other causes; the contractor shall take the necessary measures to immediately revoke access to VA network, property, information, and information systems (logical and physical) by contractor/subcontractor personnel. These measures include (but are not limited to): removing and then securing Personal Identity Verification (PIV) badges and PIV – Interoperable (PIV-I) access badges, VA-issued photo badges, credentials for VA facilities and devices, VA-issued laptops, and authentication tokens. Contractors shall notify the appropriate VA COR/CO immediately to initiate access removal.

  - Contractors/subcontractors who no longer require VA accesses will return VA-issued property to VA. This property includes (but is not limited to): documents, electronic equipment, keys, and parking passes. PIV and PIV-I access badges shall be returned to the nearest VA PIV Badge Issuance Office. Once they have had access to VA information, information systems, networks and VA property in their possessions removed, contractors shall notify the appropriate VA COR/CO.

- TRAINING.

  - All contractors and subcontractors requiring access to VA information and VA information systems shall successfully complete the following before being granted access to VA information and its systems:

    - VA Privacy and Information Security Awareness and Rules of Behavior course (Talent Management System (TMS) #10176) initially and annually thereafter.

    - Sign and acknowledge (electronically through TMS #10176) understanding of and responsibilities for compliance with the Organizational Rules of Behavior, relating to access to VA information and information systems initially and annually thereafter; and

    - Successfully complete any additional cyber security or privacy training, as required for VA personnel with equivalent information system or information access [to be defined by the VA program official and provided to the VA CO for inclusion in the solicitation document – i.e., any role-based information security training].

  - The contractor shall provide to the COR/CO a copy of the training certificates and certification of signing the Organizational Rules of Behavior for each applicable employee within five days of the initiation of the contract and annually thereafter, as required.

  - Failure to complete the mandatory annual training is grounds for suspension or termination of all physical or electronic access privileges and removal from work on the contract until such time as the required training is complete.

- SECURITY INCIDENT INVESTIGATION

  - The contractor, subcontractor, their employees, or business associates shall immediately (within one hour) report suspected security / privacy incidents to the VA OIT’s Enterprise Service Desk (ESD) by calling (855) 673-4357 (TTY: 711). The ESD is OIT’s 24/7/365 single point of contact for IT-related issues. After reporting to the ESD, the contractor, subcontractor, their employees, or business associates shall, within one hour, provide the COR/CO the incident number received from the ESD.

  - To the extent known by the contractor/subcontractor, the contractor/ subcontractor's notice to VA shall identify the information involved and the circumstances surrounding the incident, including the following:

    - The date and time (or approximation of) the Security Incident occurred.

    - The names of individuals involved (when applicable).

    - The physical and logical (if applicable) location of the incident.

    - Why the Security Incident took place (i.e., catalyst for the failure).

    - The amount of data belonging to VA believed to have been compromised.

    - The remediation measures the contractor is taking to ensure no future incidents of a similar nature.

  - After the contractor has provided the initial detailed incident summary to VA, they will continue to provide written updates on any new and relevant circumstances or facts they discover. The contractor, subcontractor, and their employes shall fully cooperate with VA or third-party entity performing an independent risk analysis on behalf of VA. Failure to cooperate may be deemed a material breach and grounds for contract termination.

  - VA IT contractors shall follow VA Handbook 6500, Risk Management Framework for VA Information Systems VA Information Security Program, and VA Information Security Knowledge Service guidance for implementing an Incident Response Plan or integrating with an existing VA implementation.

  - In instances of theft or break-in or other criminal activity, the contractor/subcontractor must concurrently report the incident to the appropriate law enforcement entity (or entities) of jurisdiction, including the VA OIG, and the VA Office of Security and Law Enforcement. The contractor, its employees, and its subcontractors and their employees shall cooperate with VA and any law enforcement authority responsible for the investigation and prosecution of any possible criminal law violation(s) associated with any incident. The contractor/subcontractor shall cooperate with VA in any civil litigation to recover VA information, obtain monetary or other compensation from a third party for damages arising from any incident, or obtain injunctive relief against any third party arising from, or related to, the incident.

  - The contractor shall comply with VA Handbook 6500.2, Management of Breaches Involving Sensitive Personal Information, which establishes the breach management policies and assigns responsibilities for the oversight, management and reporting procedures associated with managing of breaches.

  - With respect to unsecured Protected Health Information (PHI), the contractor is deemed to have discovered a data breach when the contractor knew or should have known of breach of such information. When a business associate is part of VHA contract, notification to the covered entity (VHA) shall be made in accordance with the executed BAA.

If the contractor or any of its agents fails to protect VA sensitive personal information or otherwise engages in conduct which results in a data breach involving any VA sensitive personal information the contractor/subcontractor processes or maintains under the contract; the contractor shall pay liquidated damages to the VA as set forth in clause 852.211-76, Liquidated Damages—

End of SOW