REQUIREMENT SPECIFICATIONS

FOR EQUIPMENT FOR THE AGENCY TECHNICAL SUPPORT EQUIPMENT CACHE

1.0 Scope. Occupational Safety & Heath Administration (OSHA) needs new equipment in various field offices.  OSHA Technical Center (OTC) purchases all the equipment through the Agency Technical Support Equipment Cache (TSEC).

2.0 Overview. The equipment and accessories being requested with this order is of makes and models which are currently owned by the Agency and/or can be serviced (repaired and calibrated) by OTC. The Center has the institutional knowledge and experience to service the equipment in the Agency’s best behalf including, where appropriate, having the necessary repair parts, technical drawings and schematics, calibration systems and procedures to service the technical equipment.

3.0 Requirements.

The following item requested is similar or compatible to the Larson Davis Noise Dosimeter, Part #730IS-PK05-US:

Must meet the following specifications:

- Shall include (5) personal noise dosimeters, (1) acoustical calibrator, docking/charging station to accommodate (5) noise dosimeters, microphone windscreens, and carrying case.

- Noise dosimeter must meet or exceed the performance characteristics for personal noise dosimeters specified in American National Institute (ANSI) S1.25.

- Shall be intrinsically rate for Class 1 Division 1 and/or Class 1, Zone 0.

- Shall have Bluetooth connectivity for remote communication with the noise dosimeter.

- Shall have compatible American National Institute (ANSI S1.30 compliant class 2 or better acoustical calibrator available.

- Shall have ability to download, store, format, and generate reports from measurement data using Windows 11 compatible software.

- Shall have ability to select subsets of recorded study time to include or remove from included measurement data.

- Shall allow for internal memory storage of pre and post study calibration data.  Calibration events must be included in downloaded data.

- Shall offer calibration and software support, including ability to purchase.

- Shall offer product manual.

- Vendor shall ensure the equipment manufacturer shall make replacement components available to OTC for intermediate level maintenance including, but not limited to the following components:  rechargeable batteries, microphones, windscreens, mounting clips, and microphone adapters.

- Shall make available any components needed to allow for direct signal injection into the dosimeter for testing and validation purposes.

- Vendor shall ensure the equipment manufacturer shall provide training and documentation in support of OTC’s efforts to perform in house calibration and repair of purchased equipment.  Training may be conducted in person or virtually.  Any specialized training, parts, or equipment required to perform calibration or repair of components identified as a required repair component shall be offered.

3.1 Once the equipment is received at OTC and verified, if a problem exists with the equipment the vendor will issue a return authorization and pay to have the supplies shipped back to the vendor.  The supplies will be repaired or replaced (at the discretion of the vendor) free of charge under warranty.

4.0 Delivery. Delivery will be 15 days after award and shipped through a Government Shipping Account with tracking provided.  Delivery orders will be issued and will not exceed a total of 164 kits per the specified period of performance found in the BPA.

5.0 Delivery Location

OSHA Technical Center

8660 S. Sandy Pkwy.

Sandy, UT  84070