| SUBJECT* | Non-Invasive Ventilators |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 29403-5799 |
| --- | --- |
| SOLICITATION NUMBER* | 36C24726Q0435 |
| RESPONSE DATE/TIME/ZONE | 04-09-2026 16:30 EASTERN TIME, NEW YORK, USA |
| ARCHIVE | 30 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | 6515 |
| NAICS CODE* | 339112 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Ralph H. Johnson VA Medical Center 109 Bee Street Charleston SC 29403-5799 |
| POINT OF CONTACT* | Contracting Officer Darius Crane darius.crane@va.gov (843) 577-5011 |

PLACE OF PERFORMANCE

| ADDRESS | Department of Veterans Affairs |
| --- | --- |
|  | Birmingham VAMC |
|  | VISN 07 |
|  | 700 19th St. S |
|  | Birmingham AL |
| POSTAL CODE | 35233-1927 |
| COUNTRY | USA |

ADDITIONAL INFORMATION

| AGENCY’S URL |  |
| --- | --- |
| URL DESCRIPTION |  |
| AGENCY CONTACT’S EMAIL ADDRESS |  |
| EMAIL DESCRIPTION |  |

DESCRIPTION

This Request for Information (RFI) is issued solely for information and planning purposes. It does not constitute as a solicitation.  All information received in response to this RFI that is marked as proprietary will be handled accordingly.  IAW FAR 15.201(e), responses to this notice are not offers and cannot be accepted by the Government to form a binding contract.  Responders are solely responsible for all expenses associated with responding to this RFI.

This is NOT a solicitation announcement. This is a sources sought synopsis only. This synopsis aims to gain knowledge of potential qualified sources and their size classification relative to NAICS 339112 (size standard of 1,000 employees). Responses to this synopsis will be used by the Government to make appropriate acquisition decisions. After a review of the responses to this source sought synopsis, a solicitation announcement may be published.

The Department of Veterans Affairs (VA), Regional Procurement Office-RPO East Contracting Office, is seeking sources that can provide Non-Invasive Ventilation System for the Birmingham VA Healthcare System.

The Birmingham VAMC has a requirement for 12 Non-Invasive Ventilators, to include accessories, and staff training. Ventilators should meet the following salient characteristics:

1. Patient Population & Ventilation Range

- Must support adult and pediatric patients ≥ 12.5 kg.

- Must be capable of delivering tidal volumes down to approximately 100 mL.

2. Modes of Ventilation

The ventilator shall provide the following functional therapy modes:

- Non-Invasive Ventilation (NIV) with pressure support and backup respiratory rate (e.g., Spont-PS, S/T or equivalent modes).

- Invasive Pressure-Controlled ventilation.

- Pressure-regulated volume-targeted ventilation (PRVC-type functionality).

- Proportional pressure support or equivalent capabilities to match patient effort.

- Integrated high-flow oxygen therapy capability.

3. Monitoring Capabilities

- Must provide continuous CO₂ and SpO₂ monitoring, including during non-invasive ventilation.

- Monitoring capabilities must remain available during standby to support patient readiness.

- Must store and display at least 72 hours of patient waveforms, parameters, and alarm history.

4. Portability & Power Requirements

- Turbine-driven ventilator (no external compressor required).

- Weight shall not exceed approximately 17 lbs without trolley.

- Must include:

  - Internal backup battery, and

  - Hot-swappable external battery, providing combined runtime of at least 4 hours under typical conditions.

- Must be suitable for intrahospital transport while maintaining ventilation.

5. Infection Control Features

- Must include dual HEPA filtration or equivalent capability:

  - One filter protecting external air intake.

  - One filter on the patient gas pathway/exhalation side.

6. User Interface & Display

- Minimum 12-inch color touchscreen or equivalent interface.

- Must provide intuitive graphical monitoring and indicators for alarms and patient status.

7. Connectivity

- Must support connectivity to common patient monitoring or EMR integration middleware (no brand required). Examples include:

  - Bedside monitor interface (e.g., via HL7, serial, or proprietary cable systems)

  - EMR middleware compatibility

8. Regulatory Requirement

- Must be a FDA-cleared Class II medical device meeting all applicable ventilation safety standards.

9. Accessories (Minimum Required)

- Non-invasive ventilation masks or interface.

- Breathing circuits compatible with both invasive and non-invasive modes.

- Capnography sensor/interface suitable for NIV use.

- Power supply/charger and mounting hardware.

The information identified above is intended to be descriptive, not restrictive, and to indicate the quality of the supplies/services will be satisfactory. It is the responsibility of the interested source to demonstrate to the government that the interested parties can provide the supplies/services that fulfill the required specifications. If you are interested and are capable of providing the sought-out supplies/services, please provide the requested information as well as the information indicated below.

Responses to this notice shall be submitted via email to darius.crane@va.gov no later than 16:30 EST. April 09, 202. Please include in your response company name, address, point of contact, and size of business. Include in the subject line of your email “36C24726Q0435 NON-INVASIVE VENTILATORS”. Telephone responses shall not be accepted.

If a solicitation is issued it shall be announced later, and all interested parties must respond to that solicitation announcement separately from the responses to this source sought.