- TITLE

Rees validation and calibration services for the current Rees NIH160 system.

- BACKGROUND

The DLM Branch has critical research/equipment that Rees Scientific Corp., currently monitors.  To ensure the safety and integrity of the products and research samples, and to meet CAP and cGMP requirements, equipment used in manufacturing processes must be monitored for performance to ensure the processes have not been compromised.  The sensors are monitoring laboratory equipment, including freezers, refrigerators, incubators, which contain clinical and research material critical to programs in the DLM. The primary monitored parameters are temperature, room temp and humidity.

2.1	OBJECTIVE

This request is for the validation and calibration of an existing environmental monitoring system, which is a regulatory requirement for the DLM facilities.  This service would be applicable for all existing and functioning REES inputs and associated software.  Regulatory guidance requires a fully functional, calibrated, and validated system to monitor the equipment and facility. As well as annually qualify and validate the system to ensure products are always manufactured and housed under safe and controlled conditions.

The monitors are connected by hardwired and wireless transmission, to data processing and storage equipment. Backup and security facilities are part of the monitoring system.  The data is transmitted to a standalone server.  The system maintains long-term records generated by the monitors, which provide a history of equipment performance.  Reports of monitored equipment and the facility are also produced.  Both real time and historic data are accessible in a secure fashion on workstations.

This system and the associated calibration and validation process is a propriety one that requires the service be performed by trained Rees technicians.

- CONTRACT REQUIREMENTS/ AND PERSONNEL QUALIFICATIONS

The service contract shall include all items outlined in the attached quote. This includes an annual qualification and validation of the entire system and software.  As part of the validation, a complete testing protocol that tests and reports at a minimum on the following:

  - Calibration, qualification, and validation of all exciting sensors.  This includes factory calibration at time of purchase of certified combination temperature/ humidity sensor.

  - All sensors (inputs) will be validated and calibrated to confirm these are working properly.

  - All sensors, hardwired or wireless, are installed in the correct places as indicated by their IDs and locations listed for users of their monitored equipment, in the complete Rees node program printout and online.

  - Calibration of all differential pressure sensors will be performed and within the acceptable tolerance using an NIST traceable instrument within date certificates

  - Calibration of all CO2 sensors will be performed and within the acceptable tolerance using an NIST traceable instrument within date certificates

  - Calibration of all Temperature Input of Sensors using a NIST traceable instrument will be performed and within the acceptable tolerance using an NIST traceable instrument within date certificates.  Rees system upgrade to allow for connection to an SQL database using TLS 1.2. All traffic to and from the database will be encrypted with the latest protocol, preventing eaves dropping for passwords and risk of hacking. Included with this upgrade, SNMP will be disabled in all of your V2/Z3 network devices, and the latest Rees Scientific approved version of Apache Web Server with only strong cipher suites enabled will be installed.

  - Validation Kit for New Systems and existing systems with Rees Software Ver. 3.0 Build 1020.0.24 or 1021.0.9 or higher

  - The REES system must meet 21CFR211.68(a), 21CFR211.68(b), and 21CFR11 requirements pertaining to calibration, inspection, record retention, change management, input/output verification, audit trail, data back-up, electronic signatures, electronic data, qualification, and validation.

  - At minimum, the system shall undergo defined lifecycles including, but not limited to, Design Qualification, Installation Qualification, Operational Qualification, and Performance Qualification. If data is leveraged from a prior qualification program, access to all data meeting the above stated requirements shall be requested by NIH and reviewed appropriately. All changes following NIH review shall be subjected to change control.

  - Validation Kit for Centron with Centron Presidio Software.

  - Confirm all sensors are reading correctly.

  - Confirm all sensors are properly calibrated.

  - Each piece of monitored equipment will display a certification label for each sensor. Writing on the certificate will be protected against loss of legibility following treatment of a room and equipment with cleaning and antiseptic solutions.

  - The system properly identifies an alarm situation.

  - The local alarm system is operational.

  - The remote alarm contacts are functional.

  - The proper voice and data telephone calls are made and emails sent.

  - The system accepts the proper acknowledging signal over the phone.

  - The Processor, RAM, and PROM will be tested to ensure that the program is as originally installed.

  - The A/D converter, voice chip, phone interface, display interface, and remote panel communications will be tested.

  - All output relays that are in use will be tested to ensure proper function.

  - Comprehensive Once Yearly IQ/OQ of System with full documentation provided to satisfy regulatory authorities. Comprehensive system test verifies, calibrates and documents all system functions including readings, databases, alarms, dial-outs, outputs, reports, and user programming.

  - Includes One Year Onsite Service Plan with 24/7 Technical Support.

  - Includes approximately 100 page GMP/GLP IQ/OQ validation report, 1000 page software validation test report, calibration of temperature and humidity with traceable equipment, copies of reference equipment calibration certificates and onsite refresher training at PM visit.

  - At a minimum, the vendor shall have in place a quality management system similar to or exceeding ISO 9001:2015 requirements. The vendor must provide services meeting, at minimum, the requirements of ISO/IEC 17025:2017 applicable to calibration laboratories. The vendor shall be subjected to periodic NIH audits to assess cGMP and CAP compliance to applicable quality systems. A change management process must be established by the vendor to manage all software and hardware changes to the system. All quality events shall be investigated and resolved with input from end user. All personnel performing contracted service for the NIH shall have combination of education, training, experience, and certification needed to work in a CAP and cGMP environment. Records of all personnel qualifications shall be submitted to NIH for review before personnel are allowed to perform assigned task. All subcontracting shall be pre-approved by the NIH before implementation. Calibration stickers will have to be placed on all calibrated probe/sensor.

3.1	SERVICE EXCLUSIONS

The Contractor shall not be responsible for any repairs necessitated by abuse, neglect, vandalism, fire or water.  These repairs shall be the subject of a separate purchase order and shall not be performed under this contract.

- TYPE OF ORDER

This is a severable Firm Fixed Price purchase order.

The services acquired under this contract are severable services. Funds are only available for the use of the contract line item number (CLIN) to which they are obligated. Unused funds from one CLIN or option period may not rollover for use in other periods or CLINs.

- TECHNICAL POINT OF CONTACT (TPOC)

TBD at award.

- PERIOD OF PERFORMANCE

This award shall have a base period of twelve (12) months to run 4/24/2026 thru 4/23/2027 - $21,593.43

- PLACE OF PERFORMANCE

Onsite services shall be performed at the following location:

NIH Bldg #10 - DLM

10 Center Drive

Room 2C324

Bethesda, MD 20892