BASE + 4 YEARS CERTIFICATIONS OF PHARMACY HOODS

| SUBJECT* | 575 / Base +4 / Sole Source / Yellow Bridge / 575-26-3-036-0138 (VA-26-00034672) |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 80111 |
| --- | --- |
| SOLICITATION NUMBER* | 36C25926Q0365 |
| RESPONSE DATE/TIME/ZONE | 03-31-2026 08:00am EASTERN TIME, NEW YORK, USA |
| ARCHIVE | 30 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | H366 |
| NAICS CODE* | 561720 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Network Contracting Office NCO 19 6162 South Willow Drive, Suite 300 Greenwood Village CO 80111 |
| POINT OF CONTACT* | Contracting Officer Kenneth Finchum kenneth.finchum@va.gov 303-712-5746 |

PLACE OF PERFORMANCE

| ADDRESS | Department Of Veterans Affairs |
| --- | --- |
|  | VA Western Colorado Health Care System |
|  | 2121 North Ave |
|  | Grand Junction |
|  | CO |
| POSTAL CODE | 81501 |
| COUNTRY |  |

ADDITIONAL INFORMATION

| AGENCY’S URL |  |
| --- | --- |
| URL DESCRIPTION |  |
| AGENCY CONTACT’S EMAIL ADDRESS | kenneth.finchum@va.gov |
| EMAIL DESCRIPTION | Kenneth Finchum |

DESCRIPTION

Sources Sought Notice - BASE + 4 YEARS CERTIFICATIONS OF PHARMACY HOODS

## Disclaimer:

This Request for Information (RFI) is issued solely for market research information and planning purposes and does not constitute a solicitation. All information received in response to this RFI that is marked as proprietary will be handled accordingly. Responses to this notice are NOT offers and cannot be accepted by the Government to form a binding contract. Responders are solely responsible for all expenses associated with responding to this RFI.

Information Requested from Industry:

- In response to the Sources Sought, interested Contractors shall submit the following information and any Capabilities/Qualifications Statement to include but not limited to an overview of proposed solution(s) and a description of the product your company possesses in accordance with the Statement of Work draft / Salient Characteristics List.

- Responses to market research table questionnaire shall be submitted. If this information is not provided, then it will be assumed the entire requirement cannot be met.

Description of Requirement:

The Department of Veterans Affairs, Veterans Health Administration (VHA), Network Contracting Office 19 (NCO 19) is conducting a market survey and is seeking potential sources for BASE + 4 YEARS CERTIFICATIONS OF PHARMACY HOODS for the Grand Junction VA Medical Center. This Sources Sought Notice is issued for information and planning purposes only. This is not a solicitation or a request for proposal and shall not be construed as an obligation or commitment by the Government. An award will not be made on any offers submitted in response to this notice, and it shall not be implied the Government is committed to providing any solicitation or award following this notice. The Government will not pay for any information received in response to this request, nor will the Government compensate a respondent for any costs incurred in developing the information provided.

Requirement Item

| Item No. | Part Number | Item Description |
| --- | --- | --- |
| 1 |  | BASE + 4 YEARS CERTIFICATIONS OF PHARMACY HOODS |

Required Salient Characteristics OR SOW Requirements:

SOW that the company and the technicians that come out to certify the IV Hoods MUST be CETA CERTIFIED.

Please see attached SOW

| In response to this announcement, please provide the information requested below requested in this market research table questionnaire. If this information is not provided, then it will be assumed the entire requirement cannot be met: |  |  |  |
| --- | --- | --- | --- |
| NAICS Code: |  |  |  |
| Company Name: |  |  |  |
| Address: |  |  |  |
| UEI (Unique Entity ID) Number: |  |  |  |
| Contact Name: |  |  |  |
| Phone No.: |  |  |  |
| Email: |  |  |  |
| Business Size Information - Select all that applies: - Small Business - Emerging Small Business - Small Disadvantaged Business - Certified under Section 8(a) of the Small Business Act - HUBZone - Woman Owned - Certified Service-Disabled Veteran Owned Small - Veteran Owned Small Business - Large Business |  |  |  |
| FSS/GSA Contract Holder: |  |  | Yes No |
| FSS/GSA Contract Number: |  |  |  |
| Effective Date/ Expiration Date: |  |  |  |
| Proposed solution is listed and available on the above FSS/GSA Contract: |  |  | Yes No |
| Available pricing structure of proposed solution (select all that are applicable below. |  |  |  |
| Pricing Model | Please Indicate Availability Below: (Yes / No / NA) |  |  |
|  | All on FSS | Open Market only | Mix of FSS & Open Market |
| (CPRR) Cost Per Reportable Result |  |  |  |
| Cost Per Test(CPT) |  |  |  |
| Reagent Rental Agreement |  |  |  |
| Equipment Rental with Reagent Purchase |  |  |  |
| Fixed Monthly Charge |  |  |  |
| Firm-Fixed Price |  |  |  |
| Other: (Please explain) |  |  |  |
| Federal Acquisition Regulation (FAR) Market Rearch Questions: |  |  |  |
| Buy American Act RFO 52.225-1 and RFO 52.225-2 Buy American Certificate – What percentage of the proposed product (including leases and rentals) is a: |  |  |  |
| - Domestic end product? _____________ (%) |  |  |  |
| - Foreign end product? _______________ (%) Country of Origin: |  |  |  |
| Questions for Small Businesses ONLY: |  |  |  |
| Limitations on Subcontracting RFO 52.219-14 - What percentage of the work would be subcontracted to another company? ________ - What is the business size of the subcontractor?______ Per RFO 52.219-14 the following conditions must be met: Services (except construction), it will not pay more than 50 percent of the amount paid by the Government for contract performance to subcontractors that are not similarly situated entities. Any work that a similarly situated entity further subcontracts will count towards the prime contractor's 50 percent subcontract amount that cannot be exceeded. When a contract includes both services and supplies, the 50 percent limitation shall apply only to the service portion of the contract; Supplies (other than procurement from a nonmanufacturer of such supplies), it will not pay more than 50 percent of the amount paid by the Government for contract performance, excluding the cost of materials, to subcontractors that are not similarly situated entities. Any work that a similarly situated entity further subcontracts will count towards the prime contractor's 50 percent subcontract amount that cannot be exceeded. When a contract includes both supplies and services, the 50 percent limitation shall apply only to the supply portion of the contract; |  |  |  |
| - If subcontracting, what added value do you offer (RFO 52.215-23): |  |  |  |
| - Nonmanufacturer Rule (RFO 52.219-33): - Does your company manufacturer these proposed items? [ ] yes [ ] no - Does your company exceed 500 employees? [ ] yes [ ] no If yes, list # of employees: _________ - Does your company primarily engaged in the retail or wholesale trade and normally sells the type of item being supplied? [ ] yes [ ] no - Does your company take ownership or possession of the item(s) with its personnel, equipment or facilities in a manner consistent with industry practice? [ ] yes [ ] no - Does your company supply the end item of a small business manufacturer, processor or producer made in the United States, or obtains a waiver of such requirement pursuant to paragraph (b)(5) CFR 121.406. [ ] yes [ ] no If yes, what is the manufacturer’s name? ________________ |  |  |  |

Response Instructions:

- Do not provide more than 8 pages, including cover letter page.

- Submit your response via email to kenneth.finchum@va.gov

- Submit your response NLT 08:00 (MST) on Tuesday, March 31, 2026.

- Mark your response as “Proprietary Information” if the information is considered business sensitive.

- VA has identified the appropriate North American Industry Classification System (NAICS) Code 561720 Janitorial Services Please identify and explain any NAICS codes your company believes would better represent the predominant work included in the attached Statement of Work.

(End of Sources Sought)

Statement of Work

BASE + 4 YEARS CERTIFICATIONS OF PHARMACY HOODS

- General Objectives and Requirements

  - Purpose: To ensure the comprehensive certification of all sterile compounding hoods and rooms.

  - Authority for the Contract:

    - 38 U.S. Code § 7301.b “The primary function of the Veterans Health Administration is to provide a complete medical and hospital service for the medical care and treatment of veterans, as provided in this title and in regulations prescribed by the Secretary pursuant to this title.”

    - COR/POC: Shauna Hartbauer
Email: shauna.hartbauer@va.gov
Phone: 9702420731 ext. 4426

- Background. The VA Western Colorado Health Care System, located in Grand Junction, CO, maintains segregated compounding areas (SCA)s and a containment segregated compounding area (C-SCA) which are used to aseptically compound non-hazardous and hazardous drugs, respectively.  These areas and their equipment require regular testing and certification to ensure compliance with applicable standards and to maintain a safe and controlled environment for the preparation of medications. This project involves engaging a certified contractor to furnish all necessary labor, equipment, and software to execute comprehensive testing and certification of the sterile compounding areas and associated equipment. The goal is to ensure the facility's continuous compliance with regulatory requirements and to uphold the safety and efficacy of compounded medications for patient use.

- Period of Performance.

  - Base Year: 05/01/2026 through 04/30/2027

  - Option Year 1: 05/01/2027 through 04/30/2028

  - Option Year 2: 05/01/2028 through 04/30/2029

  - Option Year 3: 05/01/2029 through 04/30/2030

  - Option Year 4: 05/01/2030 through 04/30/2031

- Hours of Operation.

  - Normal hours of operation are Monday through Friday, 0800 to 1700, except Federal Holidays listed below and any other day declared a federal holiday by the President of the United States.  The Contractor may schedule work outside of normal hours with prior approval of Contracting Officer (CO) and/or station Contracting Officer’s Representative (COR) and/or the Chief of Pharmacy.

New Year’s Day	Martin Luther King Jr’s Day	Presidents Day

Memorial Day		Independence Day		Labor Day

Columbus Day		Veterans Day			Thanksgiving Day

Christmas Day

- Scope.

  - Contractors must be certified by the Controlled Environment Testing Association (CETA).  The inspection contractor shall furnish labor, equipment, and software necessary to provide testing and certification of the Pharmacy Sterile Compounding suite and equipment.  The equipment is located at Grand Junction, CO, on the campus of the VA Western Colorado Health Care System.

  - Inspections shall include the following testing and certification of primary and secondary engineering controls:

    - Non-viable and viable testing and certification of one compounding aseptic containment isolator (CACI)

      - Serial number: CHEMOSV001-51

    - Non-viable and viable testing and certification of one compounding aseptic isolator (CAI)

      - Serial number: Micro 4x2 LF-12

    - Non-viable and viable testing and certification of one Laminar Airflow Workbench (LAFW)

      - Serial number: 4S-15-BH-19812

    - Pressure differential and air exchange testing of one containment segregated compounding area (C-SCA)

  - The contractor shall perform appropriate tests and procedures to be employed on the various primary and secondary engineering controls. The tests used will comply with current Controlled Environment Testing Association (CETA) standards, United States Pharmacopeia Chapters <797> and <800>, and current NSF standards.

- Qualifications of Field Service Engineers/Technicians. Technicians performing inspection and certification must be CETA certified.

- Response Times

  - Inspections are to be performed semi-annually (every 6 months) according to the following schedule:

    - Certification CAI - August

    - Certification C-SCA, CACI AND LAFW - November

    - Certification CAI - February

    - Certification C-SCA, CACI AND LAFW - May

  - If any of the above inspections fail certification during their scheduled semi-annual dates, the contractor will address the issues in coordination with the compounded sterile products (CSP) pharmacist.

    - The contractor will repair any mechanical issues at the time of certification if within the scope of their expertise and if all necessary parts and equipment are available.

    - If the required parts and equipment are not available at the time of certification failure, the contractor will work with the CSP pharmacist to order the needed items and return to repair and recertify the failed hood as soon as possible.

  - The Contractor (or Representative) shall contact the COR and CSP pharmacist to schedule work and prior to beginning work on site.

  - The Contractor shall conduct a joint inspection with the CSP pharmacist upon certification and testing of equipment

- Parts.

  - The contractor shall furnish all parts and consumables required to conduct testing and the price of the parts and consumables shall be included in the price of this contract.

- Service Tools and Test Equipment.

  - The contractor shall provide all tools, test equipment, manuals, and diagnostic software necessary to perform the work.

- Documentation of Service and Inspections (Quality Assurance).

  - The Contractor shall provide documentation of the services rendered for each scheduled inspection.  The documentation shall contain the following information:

    - Date of Testing

    - Name of contractor, contact information, and contract number.

    - Name and certification of technician who performed services.

    - Confirmation that the most current version of CETA Certification Guide was used.

    - Calibration certificates of the Contractor’s equipment used during the testing.

    - Contractor service ESR number/log number.

    - Identification of equipment serviced: Inventory ID number, Manufacturer's name, device name, model number & serial number, any other manufacturer's identification numbers.

    - Itemized Description of Service Performed (including, if applicable, costs associated with after-normal-working-hour services) including Labor and Travel, Parts (with part numbers), Materials and Circuit Location of problem/corrective action.

    - Location of the data collected.

    - Explanation of test procedures used and justification for any deviations from established industry practices encountered during the certification process.

    - Comparison of collected data with expected values or specific equipment performance criteria.

    - Recommendations for repairs or improvements.

    - Signatures of the FSE/T performing services described.

- List of Testing and Certification Requirements.

  - Contractor shall describe proposed tests for each listed area including but not restricted to the tests listed here, based on Contractor’s expertise in the applicable standards.

    - Chemo preparation area (637 cubic feet).  Pressurization to be negative to Anteroom.  Supply Air:760 CFM; Exhaust Air:900 CFM; Outside Air:380 CFM; Return Air: 0 CFM; Flow Differential: -140 CFM; minimum air changes per hour: 30; Minimum pressure differential -0.01 inch to -0.03 inch.

      - Test Room Airflow and Pressure

      - Air Changes per Hour

    - Compounding Aseptic Containment Isolator (Chemosphere: Isotech 19CS07P14)

      - Test Airflow and Pressure

      - Air Changes per Hour

      - Test Laminar Air Flow (Laminar Airflow visualization study)

      - Test HEPA Filter Integrity

      - Calibrate pressure gauge

      - Viable Airborne Particle Sample, with Colony ID

      - Viable Surface Particle Sample, with Colony ID

      - Environmental non-viable particle count

    - Laminar Airflow Workbench (LAFW) Serial number: 4S-15-BH-19812

      - Test Airflow and Pressure

      - Air Changes per Hour

      - Test Laminar Air Flow (Laminar Airflow visualization study)

      - Test HEPA Filter Integrity

      - Calibrate pressure gauge

      - Viable Airborne Particle Sample, with Colony ID

      - Viable Surface Particle Sample, with Colony ID

      - Environmental non-viable particle count

    - Compounding Aseptic Isolator (Microsphere: Isotech 015G07P34)

      - Test Airflow and Pressure

      - Air Changes per Hour

      - Test Laminar Air Flow (Laminar Airflow visualization study)

      - Test HEPA Filter Integrity

      - Calibrate pressure gauge

      - Viable Airborne Particle Sample, with Colony ID

      - Viable Surface Particle Sample, with Colony ID

      - Environmental non-viable particle count

- Risk Control.  The Contractor shall be escorted by the Agency. The Contractor may need passes from the on-station Police Department.

- Invoices and Payment. Payment will be made in arrears of each inspection. Invoices must include, as a minimum, the following information: Contract Number, Purchase Order Number, Equipment covered (to include serial #s) and date of service. Any additional charges must be approved by the Contracting Officer before service is initiated.

- Safety Requirements. The Contractor is solely responsible for compliance with all federal, state and local laws, the Occupational, Safety and Health Act (OSHA) (Public Law 91-596) and the resulting standards, OSHA Standards 29 CFR 1910 and 1926, as applicable, and the protection of their employees and subcontractor employees.

- VA Regulations. Contractor employees always wear visible identification when on Station.  Consumption of tobacco and other substances legal for use in the State of Colorado is not permitted on VA property, in any form.  Possession of firearms is prohibited.  Containers, including personal bags and toolboxes, are subject to search at the discretion of the VA Police.