| SUBJECT* | PHACOEMULSIFICATION UNITS |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 98661 |
| --- | --- |
| SOLICITATION NUMBER* | 36C26026Q0530 |
| RESPONSE DATE/TIME/ZONE | 04-20-2026 6:00 PACIFIC TIME, LOS ANGELES, USA |
| ARCHIVE | 30 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | 6540 |
| NAICS CODE* | 339112 |
| CONTRACTING OFFICE ADDRESS | Network Contracting Office 20 ATTN: V4CONT 1601 E. Fourth Plain Blvd Bldg.17, Suite B428 Vancouver WA 98661 |
| POINT OF CONTACT* | Contract Specialist Gregory Watson Gregory.Watson2@va.gov 360-553-7602 |

PLACE OF PERFORMANCE

| ADDRESS | Department of Veterans Affairs |
| --- | --- |
|  | Portland VA Medical Center |
|  | 3710 SW US Veterans Hospital Rd. |
|  | Portland OR |
| POSTAL CODE | 97239 |
| COUNTRY | USA |

ADDITIONAL INFORMATION

| AGENCY’S URL | https://www.va.gov/portland-health-care/ |
| --- | --- |
| URL DESCRIPTION | https://www.va.gov/portland-health-care/ |
| AGENCY CONTACT’S EMAIL ADDRESS | Gregory.Watson2@va.gov |
| EMAIL DESCRIPTION | Contract Specialist |

DESCRIPTION

THIS IS A REQUEST FOR INFORMATION ONLY. The purpose of this notice is to obtain information regarding the availability and capability of all qualified sources to fulfil a potential requirement.  The responses received from interested contractors will assist the Government in determining the appropriate acquisition method.

The Department of Veterans Affairs (VA), Network Contracting Office (NCO) 20, is conducting market research to identify potential sources for:

Phacoemulsification Units for the Portland VAMC. The Phacoemulsification Units are brand name or equal to Phacoemulsification Units manufactured by Alcon Vision LLC.

| Quantity | Product Code | Description |
| --- | --- | --- |
| 30 | 8065000315 | Unity 4D Phaco Handpiece |
| 3 | 8065000215 | OZIL FRAG HANDPIECE |
| 3 | 900484665 | UNITY VCS VLTP |
| 1 | 8065753167 | ASSY,SHIP,VIT 23GA ANT |
| 1 | 8065000339 | ANTERIOR PAK UCS .9ULTRA 30BAL |
| 1 | 8065000393 | NG VISCOUS FLUID CONTROL |
| 1 | 8065000329 | POSTERIOR PAK UVRC 23G BEV 10K S |
| 1 | 8065797102 | ISPAN C3F8 PERFLUOROPROPANE |
| 1 | 8065797002 | ISPAN SF6 SULFUR HEXAFLUORIDE 125GM |

Salient Characteristics Phacoemulsification and Vitreoretinal Surgical System:

Vitreoretinal Capabilities (Minimum Requirements)

• Provides a vitrectomy cutter capable of at least 30,000 cuts per minute for efficient vitreous removal.

• Supports compatibility with dual-blade or equivalent high efficiency cutters for close proximity retinal work.

• Maintains stable intraocular pressure (IOP) through real time pressure sensing with independent control of infusion and aspiration, accommodating the range of viscosities encountered in vitreoretinal surgery.

• Provides aspiration modes that allow transition between flow-based and vacuum-based aspiration to support both shave vitrectomy and posterior vitreous detachment techniques.

• Includes technology that detects aspiration occlusion or changes in viscosity and adjusts aspiration rate accordingly to enhance safety and efficiency, including during training scenarios.

• Offers RGB or equivalent multi channel LED illumination with tunable intensity and color profiles, along with independent control for multiple instruments.

• Provides illumination with customizable spectral output to reduce high wavelength exposure (blue/green) and align with aphakic hazard safety profiles.

• Supports compatibility with 27 gauge vitreoretinal instruments.

• Includes an internal 532 nm laser for retinal photocoagulation.

• Supports multi spot laser capability to reduce treatment time, offering selectable spot patterns such as 4 spot, 2 spot, and 1 spot.

• Provides an entry and infusion system designed to reduce trocar insertion force and allow efficient infusion, with compatibility for 3 mm, 4 mm, and 6 mm cannulas.

• Enables rapid setup and teardown, including single step priming and testing or equivalent, with minimal required VR connections.

Cataract Surgery Capabilities (Minimum Requirements)

• Provides intelligent or dynamic fluidics that maintain chamber stability at physiologic IOP levels.

• Includes surge mitigation functionality using pressure sensing and automatic venting to reduce post occlusion surge.

• Includes a torsional or equivalent advanced phaco handpiece intended for efficient lens removal.

• Provides real time thermal sensing at or near the phaco tip with automatic power modulation to prevent corneal wound burns.

Potential candidates having the capabilities necessary to provide the above stated requirement are invited to respond to this Request for Information via e-mail to Gregory Watson at Gregory,Watson2@va.gov  no later than 20 April 2026, 06:00 PST. No telephone inquiries will be accepted.

RESPONSES SHOULD INCLUDE THE FOLLOWING INFORMATION: company name, address, and business size; point of contact name, phone number, e-mail address. NAICS Code 339112 is applicable to determine business size standard. In addition please enclose the country of origin for any items not manufactured in the USA.   Any questions or concerns may also be directed to Gregory Watson via e-mail.