Statement of Work (SOW)

- General:   The contractor shall provide full-service support for both hardware and software for the Omnicell systems located at the Eastern Colorado Healthcare System which includes the main Rocky Mountain VA Medical Center and surrounding Community Based Outpatient Clinics (CBOCs). Omnicell is the manufacturer of the pharmacy automation equipment.  Omnicell has proprietary rights to the equipment and the repair of the equipment. Due to the complexity of the equipment, it must be maintained by factory trained staff and validated prior to use.  Omnicell can perform maintenance to the required FDA and safety standards.

Description of Work: The contractor shall provide all required services and subscription to cloud based software including emergency service, repairs, preventative and corrective maintenance, labor, travel, unlimited parts, technical support, application support, and factory authorized updates/modifications/upgrades for the Omnicell systems. All services performed shall be in accordance with the original equipment manufacturer (OEM) specifications. Field service engineers must be licensed or contracted with the original manufacturer in order to work on the equipment. All parts and software must be certified new from the original manufacturer. Reference below for equipment type and the required service:

Summary of Service & Equipment Description: Service Coverage for Omnicell hardware and software are for but not limited to the model names and product numbers and coverage verbiage: SEE ATTACHED EXCEL

Hours of Performance:  Normal hours of coverage are Monday through Friday 6:00 AM to 6:00 PM MST excluding federal holidays.

- All service/repairs will be performed during normal hours of coverage unless requested or approved by the Point of Contact (POC).

- Preventative periodic maintenance will be performed Monday through Friday 6:00 AM to 6:00 PM MST or as agreed to by the user and the POC in accordance with manufacturer’s recommendations.

- There will be 24/7 telephone support.

- Federal Holidays observed by the VAMC are:

    - New Year’s Day

    - Martin Luther King Day

    - Presidents’ Day

    - Memorial Day

    - Juneteenth Day

    - Independence Day

    - Labor Day

    - Columbus/Indigenous People’s Day

    - Veterans’ Day

    - Thanksgiving Day

    - Christmas Day

NOTE: Hardware/software update/upgrade installations will be scheduled and performed at a mutually agreed upon time at no additional charge to the Government (unless it would be detrimental to equipment up-time; to be determined by the POC or designated representatives. Government provides software/hardware upgrade/update. There shall be no additional charge for time spent at the site during or after the normal hours of coverage awaiting the arrival of additional FSE and/or delivery of parts.

Conformance Standards: Contractor shall provide service and ensure that the equipment functions in accordance with the equipment manufacturer’s technical performance specifications.

Contractor shall conform to all regulations to include but not be limited to federal, state, and local governing any chemicals, equipment or work which may be used or performed in

the work under this contract.

Operational Uptime – Requirements: The medical equipment/system

shall be operable and available for use 98% of the normal operating hours of the equipment. Downtime will be computed from notification of problem during normal work hours.

Guaranteed Response Time: The OEM certified Field Service Engineer (FSE) will use commercially reasonable efforts to respond to Customer’s requests within 24 hours of notification to the contractor’s service support team.

Live Phone Support: Service shall include phone support with contractor technical support engineers 24 hours, 7 days a week.

After Hours Phone Support: Service for after hour phone support shall include 24 hours, 7 days a week, and 365 days a year for technical assistance when required.

Software Enhancements: Periodic software enhancements shall be provided as part of the firm fixed price contract and shall be considered within scope.

Preventative Maintenance: Perform system scheduled maintenance inspections as outlined in the user’s manual developed by the manufacturer.

Parts Exchange: All parts required for the repair of the system (excludes instruments and accessories) are covered.

Labor & Travel Expenses: All labor and travel conducted during normal business hours: Monday through Friday, 6:00 AM to 6:00PM MST (excludes contractor company holidays) are covered.

Certified Trained Technicians and Key Personnel: The scheduled maintenance and service calls shall be performed by OEM trained personnel in accordance with the agreement. Technicians shall include fully qualified Facility Service Engineers (FSE) and a fully qualified FSE who will serve as the backup and show proof of competency, as shown by training conducted on said system. "Fully Qualified" is based upon training and on experience in the field. For training, the FSE(s) has successfully completed a formalized training program, for the equipment identified in this statement of work. For field experience, the FSE(s) has a minimum of two years of experience (except for equipment newly on the market) with respect to scheduled and unscheduled preventive and remedial maintenance. Proof of training shall be provided upon demand and be immediately sent via fax, upon demand to the POC. All FSE’s and technicians shall be authorized by the Contractor to perform the maintenance services. All work shall be performed by "Fully Qualified" competent FSE's. The POC and/or designated representative specifically reserve the right to reject any of the Contractor's personnel and refuse them permission to work on the VAMC equipment. It is anticipated that any hospital regulatory inspections (TJC, etc.) will have the contractor providing proof that they meet industry standards of quality, and traceable standards, as used to calibrate this device- system, at the time of the maintenance event.

Period of Performance:

Base year:      05/01/2026 – 03/30/2027

Option year 1: 05/01/2027 – 03/30/2028

Dates as specified unless contract award passes these dates. Then the first date of performance will start the day of award.

Scheduled Maintenance: All labor and travel conducted during normal business hours: Monday through Friday, 8:00 AM to 6:00PM MST (excludes contractor company holidays) are covered.

Parts: All labor and travel conducted during normal business hours: Monday through Friday, 8:00 AM to 6:00PM MST (excludes contractor company holidays) are covered.

Parts and Services Not Covered or Excluded by Contract: All labor and travel conducted during normal business hours: Monday through Friday, 6:00 AM to 6:00PM MST (excludes contractor company holidays) are covered.

Service Manuals/Tools/Equipment: All labor and travel conducted during normal business hours: Monday through Friday, 6:00 AM to 6:00PM MST (excludes contractor company holidays) are covered.

Check in Requirements: The Field Service Engineer must report to the Biomedical Engineering Department to obtain a badge and sign in with the Biomedical Engineering Service before work begins.  Submit any mobile media devices that would be used on the system to a virus scan.  Upon completion of work, the Field Service Engineer must report to the Biomedical Engineering Service to sign out and brief Biomedical Staff or Supervisor if Biomedical Staff is unavailable concerning completion of service.  At the end of briefing Field Service Engineer will return the badge and sign out in the Biomedical Engineering Department.

Documentation:  Contractor shall furnish a detailed field service report upon completion of work to Biomedical Staff in the Biomed Department. Payment will not be processed until a properly completed service report is received.  The service report shall contain, at a minimum, the following information:

- Identification of equipment to be serviced:

  - VA EE Number

  - Contractor Inventory ID number

  - Manufacturer Name

  - Device Name

  - Model Number

  - Serial Number

  - System ID

- Name of contractor and contract number

- Name of FSE who performed services

- Contractor service ESR number/log number

- Date, time (starting and ending), equipment downtime and hours on-site for service call

- VA purchase order number(s) covering the call if outside normal working hours

- Total time spent performing maintenance.

- Detailed narrative and itemized description of the services required and performed including:

  - Labor and travel

  - Parts (with part numbers)

  - Materials

  - Location

  - Corrective Action

- Copies of all test reports

- Complete list of parts replaced (when applicable)

- Date and time the repair/support was completed

- Total Cost to be billed (if applicable – i.e., part(s) not covered, or service rendered after normal hours of coverage).

- Signatures:

  - FSE performing services described

  - Authorized VA employee who witnessed service described

  - NOTE: - Any additional charges claimed must be approved by the POC or designated representative before service is completed!

The service report shall itemize every item in the specification.  Each item shall state the "as found" condition or values, the "calibrated to" or "adjusted to" values, the factory design tolerances, and a complete description of all work performed concerning the items.  Included will be a list of new parts used and recommended future repairs.

Reporting Requirements: The Contractor shall be required to report to the Clinical

Engineering Department to log in. This check- in is mandatory. When the service is completed, the FSE shall document services rendered on a legible ESR(s). The FSE shall be required to log out with Biomedical Engineering and submit the ESR(s) to the POC or designated representative. ALL ESRs shall be submitted to the equipment user for an "acceptance signature" and to the POC or designated representative for an "authorization signature". If POC or designated representative is unavailable, a signed, authorized copy of the ESR will be left with the user.

Place of Performance:

Department of Veterans Affairs

Eastern Colorado VA Healthcare System

Rocky Mountain VA Medical Center

1700 N Wheeling Street

Aurora, CO 80045

Information Technology Security Requirements: The contractor, their personnel, and their subcontractors shall be subject to the Federal laws, regulations, standards, and VA Directives and Handbooks regarding information and information system security as delineated in this contract. The contractor shall comply with all Federal laws and regulations the VA has developed when VA sensitive information is accessed, used, stored, generated, transmitted, or exchanged by and between VA and a contractor. The information made available to the contractor by VA for the performance of this contract will be used only for the purposes of performance under this contract.   The certification and accreditation requirements do not apply to this requirement, and a security accreditation package is not required.

Security Statement:  Sensitive VA information is contained within the systems covered by this contract.

Biomedical Engineering shall perform virus scans on all removable media prior to use on VA medical equipment. This includes all types of removable media, including media (e.g., USB devices, CDs, dongles, etc.) that have been issued by VA, media not issued by VA, and media brought in by contractor or independent service organizations. Within accordance of VA Directive 6500, Information Security Program, September 2007

The contractor will not transfer any VA information to a location outside the VA and only to VA locations determined by the VA System Administrator.  The information in these systems may be covered by the Privacy Act 1974 which contains criminal penalties of abuse of information.

During onsite service, the contractor shall be chaperoned by VA Personnel.  However, the contractor shall not be issued a User ID/Password.

Non-volatile memory devices, working or non-working, shall NOT be removed from the Rocky Mountain Regional VA Medical Center until the ISO has certified that the data has been destroyed.   For magnetic devices and media, the data destruction will be degassing.  Other forms of cleansing will be used for non-magnetic media.

The contractor will not have remote access to complete the repair(s) and preventive maintenance.

GENERAL RULES OF BEHAVIOR

a. Rules of Behavior are part of a comprehensive program to provide complete

information security. These rules establish standards of behavior in recognition of the fact that

knowledgeable users are the foundation of a successful security program. Users must

understand that taking personal responsibility for the security of their computer and the

information it contains is an essential part of their job.

b. The following rules apply to all VA contractors. I agree to:

(1) Follow established procedures for requesting, accessing, and closing user accounts

and access. I will not request or obtain access beyond what is normally granted to users or

by what is outlined in the contract.

(2) Use only systems, software, databases, and data which I am authorized to use,

including any copyright restrictions.

(3) I will not use other equipment (OE) (non-contractor owned) for the storage, transfer,

or processing of VA sensitive information without a VA CIO approved waiver, unless it has

been reviewed and approved by local management and is included in the language of the

contract. If authorized to use OE IT equipment, I must ensure that the system meets all

applicable 6500 Handbook requirements for OE.

(4) Not use my position of trust and access rights to exploit system controls or access

information for any reason other than in the performance of the contract.

(5) Not attempt to override or disable security, technical, or management controls

unless expressly permitted to do so as an explicit requirement under the contract or at the

direction of the COTR or ISO. If I am allowed or required to have a local administrator

account on a government-owned computer, that local administrative account does not confer

me unrestricted access or use, nor the authority to bypass security or other controls

except as expressly permitted by the VA CIO or CIO's designee.

(6) Contractors’ use of systems, information, or sites is strictly limited to fulfill the terms

of the contract. I understand no personal use is authorized. I will only use other Federal

government information systems as expressly authorized by the terms of those systems. I

accept that the restrictions under ethics regulations and criminal law still apply.

(7) Grant access to systems and information only to those who have an official need to

know.

(8) Protect passwords from access by other individuals.

(9) Create and change passwords in accordance with VA Handbook 6500 on systems

and any devices protecting VA information as well as the rules of behavior and

security settings for the particular system in question.

(10) Protect information and systems from unauthorized disclosure, use, modification, or

destruction. I will only use encryption that is FIPS 140-2 validated to safeguard VA sensitive

information, both safeguarding VA sensitive information in storage and in transit regarding

my access to and use of any information assets or resources associated with my

performance of services under the contract terms with the VA.

(11) Follow VA Handbook 6500.1, Electronic Media Sanitization to protect VA

information. I will contact the COTR for policies and guidance on complying with this

requirement and will follow the COTR's orders.

(12) Ensure that the COTR has previously approved VA information for public

dissemination, including e-mail communications outside of the VA as appropriate. I will not

make any unauthorized disclosure of any VA sensitive information through the use of any

means of communication including but not limited to e-mail, instant messaging, online chat,

and web bulletin boards or logs.

(13) Not host, set up, administer, or run an Internet server related to my access to and

use of any information assets or resources associated with my performance of services

under the contract terms with the VA unless explicitly authorized under the contract or in

writing by the COTR.

(14) Protect government property from theft, destruction, or misuse. I will follow VA

directives and handbooks on handling Federal government IT equipment, information, and

systems. I will not take VA sensitive information from the workplace without authorization

from the COTR.

(15) Only use anti-virus software, antispyware, and firewall/intrusion detection software

authorized by VA. I will contact the COTR for policies and guidance on complying with this

requirement and will follow the COTR's orders regarding my access to and use of any

information assets or resources associated with my performance of services under the

contract terms with VA.

(16) Not disable or degrade the standard anti-virus software, antispyware, and/or

firewall/intrusion detection software on the computer I use to access and use information

assets or resources associated with my performance of services under the contract terms

with VA. I will report anti-virus, antispyware, firewall or intrusion detection software errors, or

significant alert messages to the COTR.

(17) Understand that restoration of service of any VA system is a concern of all users of

the system.

(18) Complete required information security and privacy training, and complete required

training for the particular systems to which I require access.

PRIVACY

- Contractors and any subcontractors must adhere to the provisions of Public Law 104-191, Health Insurance Portability and Accountability Act (HIPAA) of 1996.  This includes both the Privacy and Security Rules published by the Department of Health and Human Services (HHS).

- As required by HIPAA, HHS has promulgated rules governing the use and disclosure of protected health information by covered entities, Veterans Health Administration (VHA). In accordance with HIPAA, the contractor may be required to enter into a Business Associate Agreement (BAA) with VHA.

- Business associates must follow VHA privacy policies and practices when applicable. All contractors and business associates must receive privacy training annually.

- For contractors and business associates who do not have access to VHA computer systems, this requirement is met by completing VHA National Privacy Policy training, other VHA approved privacy training or contractor furnished training that meets the requirements of the HHS Standards for Privacy of Individually Identifiable Health Information as determined by VHA.

- For contractors and business associates who are granted access to VHA computer systems, this requirement is met by completing VHA National Privacy Policy training or other VHA approved privacy training.  Proof of training is required upon request.

Safety Requirements:  In the performance of this contract, the Contractor shall take such safety precautions as the Contracting Officer may determine to be reasonably necessary to protect the lives and health of occupants of the building.  The Contracting Officer shall notify the Contractor of any safety issues and the action necessary to correct these issues.  Such notice, when served to the Contractor or his representative at the work site shall be deemed sufficient for the corrective actions to be taken.  If the Contractor fails or refuses to comply promptly, the Contracting Officer may issue an order stopping all or part of the work and hold the Contractor in default.

Invoicing:  Payment to be made monthly in arrears by certified invoices and must contain the contract number in addition to the requirements detailed in 52.212-4 (G) to be considered valid.  All invoices shall be submitted to the VA Financial Service Center and emailed to the COTR.  All invoices will reference the purchase order number assigned to the contract.

Additional Charges: There shall be no additional charge for time spent at the site (during, or after the normal hours of coverage) awaiting the arrival of additional FSE and/or delivery of parts.

Reporting Required Services Beyond the Contract Scope: The Contractor shall immediately, but no later than 24 (twenty-four) consecutive hours after discovery, notify the POC and/or designated representative (in writing) of the existence or the development of any defects in, or repairs required, to the scheduled equipment which the Contractor considers he/she is not responsible for under the terms of the contract. The Contractor shall furnish to the POC or designated representative a written estimate of the cost to make necessary repairs.

Condition of Equipment: The Contractor shall accept responsibility for the equipment.

Failure to inspect the equipment prior to contract award will not relieve the Contractor from

performance of the requirements of this contract.

Test Equipment: Prior to commencement of work on this contract, the Contractor shall

provide the VAMC with a copy of the current calibration certification of all test equipment which is to be used by the Contractor on VAMC's equipment. This certification shall also be provided on a periodic basis when requested by the VAMC. Test equipment calibration shall be traceable to the Original Equipment Manufacture’s standard.

Identification, Parking, Smoking, and VA Regulations: The Contractor's FSE's

shall always wear visible identification while on the premises of the VAMC. The Contractor shall park in the appropriate designated parking areas. Information on parking is available from the VA Police Section. The VAMC will not invalidate or make reimbursement for parking violations of the Contractor under any conditions. Smoking is prohibited inside any buildings at the VAMC. Possession of weapons is prohibited. Enclosed containers, including tool kits, shall be subject to search. Violations of VA regulations may result in citation answerable in the United States (Federal) District Court, not a local district, state, or municipal court.

Training: The Contractor shall provide first call Clinical Engineering training for current

qualified technicians at the VAMC, CT during regular service hours of Monday through Friday from 8:00am – 6:00 PM MST. This is not an ongoing training and will only be provided on an annual basis as required. All training for upgrades to equipment shall be considered within scope throughout the life of the contract. All service and repairs are performed by OEM factory certified technicians.

Transportation Charges: All costs associated with transportation, from shipping point to

the destination specified above, and all costs of removal after contract completion, shall be paid by the Contractor and incorporated in the price proposal. The Contractor shall be responsible for all damage in transit including any transportation costs for replacement.

Compliance with OSHA Bloodborne Pathogens Standard: The Contractor shall comply with the Federal/California OSHA Bloodborne Pathogens Standard.

The Contractor shall:

- Have methods by which all employees are educated as to risks associated with

- bloodborne pathogens.

- Have policies and procedures which reduce the risk of employee exposure to bloodborne pathogens.

- Have mechanisms for employee counseling and treatment following exposure to bloodborne pathogens.

- Provide appropriate personal protective equipment/clothing such as gloves, gowns, masks, protective eyewear, mouthpieces for the employee during performance of the contract.

Miscellaneous: Additional equipment may be added and/or removed throughout the life of this contract and shall be considered within scope.

Contractor Personnel Security Requirements: Security Requirements: All contractor personnel shall obtain a short-term identification badge issued by the POC or government designated representative. Such badge shall be worn by the individual and prominently always displayed while on VA Property. No employee of the contractor shall enter the project site without a valid identification badge issued by the VA. To obtain a short-term identification badge, contractor personnel shall present to the POC a valid (non-expired) photo identification issued by a US federal, state, or local government agency. Escort will be provided as required in sensitive work areas. The C&A requirements do not apply, and the security accreditation package is not required. The identified Sole Source contractor shall be subject to the Federal laws, regulations, standards, and VA Directives and Handbooks regarding information and information system security as delineated in standard acquisition guidelines. Per the VA handbook 6500.6, no VA sensitive information will be transferred, shared, or stored with the contractor. As such, no additional mechanisms will be needed to ensure the protection of information.

Records Management Statement:

1.  Contractor shall comply with all applicable records management laws and regulations, as well as National Archives and Records Administration (NARA) records policies, including but not limited to the Federal Records Act (44 U.S.C. chs. 21, 29, 31, 33), NARA regulations at 36 CFR Chapter XII Subchapter B, and those policies associated with the safeguarding of records covered by the Privacy Act of 1974 (5 U.S.C. 552a). These policies include the preservation of all records, regardless of form or characteristics, mode of transmission, or state of completion.

2.  In accordance with 36 CFR 1222.32, all data created for Government use and delivered to, or falling under the legal control of, the Government are Federal records subject to the provisions of 44 U.S.C. chapters 21, 29, 31, and 33, the Freedom of Information Act (FOIA) (5 U.S.C. 552), as amended, and the Privacy Act of 1974 (5 U.S.C. 552a), as amended and must be managed and scheduled for disposition only as permitted by statute or regulation.

3.  In accordance with 36 CFR 1222.32, Contractor shall maintain all records created for Government use or created in the course of performing the contract and/or delivered to, or under the legal control of the Government and must be managed in accordance with Federal law. Electronic records and associated metadata must be accompanied by sufficient technical documentation to permit understanding and use of the records and data.

4. Rocky Mountain Regional VA Medical Center and its contractors are responsible for preventing the alienation or unauthorized destruction of records, including all forms of mutilation. Records may not be removed from the legal custody of the Rocky Mountain Regional VA Medical Center or destroyed except for in accordance with the provisions of the agency records schedules and with the written concurrence of the Head of the Contracting Activity. Willful and unlawful destruction, damage or alienation of Federal records is subject to the fines and penalties imposed by 18 U.S.C. 2701.

In the event of any unlawful or accidental removal, defacing, alteration, or destruction of records, Contractor must report to Rocky Mountain Regional VA Medical Center. The agency must report promptly to NARA in accordance with 36 CFR 1230.

5.  The Contractor shall immediately notify the appropriate Contracting Officer upon discovery of any inadvertent or unauthorized disclosures of information, data, documentary materials, records or equipment. Disclosure of non-public information is limited to authorized personnel with a need-to-know as described in the [contract vehicle]. The Contractor shall ensure that the appropriate personnel, administrative, technical, and physical safeguards are established to ensure the security and confidentiality of this information, data, documentary material, records and/or equipment is properly protected. The Contractor shall not remove material from Government facilities or systems, or facilities or systems operated or maintained on the Government’s behalf, without the express written permission of the Head of the Contracting Activity. When information, data, documentary material, records and/or equipment is no longer required, it shall be returned to Rocky Mountain Regional VA Medical Center control or the Contractor must hold it until otherwise directed. Items returned to the Government shall be hand carried, mailed, emailed, or securely electronically transmitted to the Contracting Officer or address prescribed in the [contract vehicle]. Destruction of records is EXPRESSLY PROHIBITED unless in accordance with Paragraph (4).

6.  The Contractor is required to obtain the Contracting Officer's approval prior to engaging in any contractual relationship (sub-contractor) in support of this contract requiring the disclosure of information, documentary material and/or records generated under, or relating to, contracts. The Contractor (and any sub-contractor) is required to abide by Government and Rocky Mountain Regional VA Medical Center guidance for protecting sensitive, proprietary information, classified, and controlled unclassified information.

7.  The Contractor shall only use Government IT equipment for purposes specifically tied to or authorized by the contract and in accordance with Rocky Mountain Regional VA Medical Center policy.

8.  The Contractor shall not create or maintain any records containing any non-public Rocky Mountain Regional VA Medical Center information that are not specifically tied to or authorized by the contract.

9.  The Contractor shall not retain, use, sell, or disseminate copies of any deliverable that contains information covered by the Privacy Act of 1974 or that which is generally protected from public disclosure by an exemption to the Freedom of Information Act.

10.  The Rocky Mountain Regional VA Medical Center owns the rights to all data and records produced as part of this contract. All deliverables under the contract are the property of the U.S. Government for which Rocky Mountain Regional VA Medical Center shall have unlimited rights to use, dispose of, or disclose such data contained therein as it determines to be in the public interest. Any Contractor rights in the data or deliverables must be identified as required by RFO 52.227-11 through RFO 52.227-20.

11.  Training.  All Contractor employees assigned to this contract who create, work with, or otherwise handle records are required to take VHA-provided records management training, Talent Management System (TMS) Item #3873736, Records Management for Records Officers and Liaisons. The Contractor is responsible for confirming training has been completed according to agency policies, including initial training and any annual or refresher training.

# APPENDIX C — VA INFORMATION AND INFORMATION SYSTEM SECURITY AND PRIVACY LANGUAGE FOR INCLUSION IN CONTRACTS, AS APPROPRIATE

- GENERAL. This entire section applies to all acquisitions requiring any Information Security and Privacy language. Contractors, contractor personnel, subcontractors and subcontractor personnel will be subject to the same federal laws, regulations, standards, VA directives and handbooks, as VA personnel regarding information and information system security and privacy.

- VA INFORMATION CUSTODIAL LANGUAGE. This entire section applies to all acquisitions requiring any Information Security and Privacy language.

  - The Government shall receive unlimited rights to data/intellectual property first produced and delivered in the performance of this contract or order (hereinafter “contract”) unless expressly stated otherwise in this contract. This includes all rights to source code and all documentation created in support thereof. The primary clause used to define Government and Contractor data rights is RFO 52.227-14 Rights in Data – General. The primary clause used to define computer software license (not data/intellectual property first produced under this contractor or order) is RFO 52.227-20.

  - Information made available to the contractor by VA for the performance or administration of this contract will be used only for the purposes specified in the service agreement, SOW, PWS, PD, and/or contract. The contractor shall not use VA information in any other manner without prior written approval from a VA Contracting Officer (CO). The primary clause used to define Government and Contractor data rights is RFO 52.227-14.

  - VA information will not be co-mingled with any other data on the contractor’s information systems or media storage systems. The contractor shall ensure compliance with Federal and VA requirements related to data protection, data encryption, physical data segregation, logical data segregation, classification requirements and media sanitization.

  - VA reserves the right to conduct scheduled or unscheduled audits, assessments, or investigations of contractor Information Technology (IT) resources to ensure information security is compliant with Federal and VA requirements. The contractor shall provide all necessary access to records (including electronic and documentary materials related to the contracts and subcontracts) and support (including access to contractor and subcontractor staff associated with the contract) to VA, VA's Office Inspector General (OIG), and/or Government Accountability Office (GAO) staff during periodic control assessments, audits, or investigations.

- The contractor may only use VA information within the terms of the contract and applicable Federal law, regulations, and VA policies. If new Federal information security laws, regulations or VA policies become applicable after execution of the contract, the parties agree to negotiate contract modification and adjustment necessary to implement the new laws, regulations, and/or policies.

- The contractor shall not make copies of VA information except as specifically authorized and necessary to perform the terms of the contract. If copies are made for restoration purposes, after the restoration is complete, the copies shall be destroyed in accordance with VA Directive 6500, VA Cybersecurity Program and VA Information Security Knowledge Service.

- If a Veterans Health Administration (VHA) contract is terminated for default or cause with a business associate, the related local Business Associate Agreement (BAA) shall also be terminated and actions taken in accordance with VHA Directive 1605.05, Business Associate Agreements. If there is an executed national BAA associated with the contract, VA will determine what actions are appropriate and notify the contactor.

- The contractor shall store and transmit VA sensitive information in an encrypted form, using VA-approved encryption tools which are, at a minimum, Federal Information Processing Standards (FIPS) 140-2, Security Requirements for Cryptographic Modules (or its successor) validated and in conformance with VA Information Security Knowledge Service requirements. The contractor shall transmit VA sensitive information using VA approved Transport Layer Security (TLS) configured with FIPS based cipher suites in conformance with National Institute of Standards and Technology (NIST) 800-52, Guidelines for the Selection, Configuration and Use of Transport Layer Security (TLS) Implementations.

- The contractor’s firewall and web services security controls, as applicable, shall meet or exceed VA’s minimum requirements.

- Except for uses and disclosures of VA information authorized by this contract for performance of the contract, the contractor may use and disclose VA information only in two situations: (i) in response to a qualifying order of a court of competent jurisdiction after notification to VA CO (ii) with written approval from the VA CO. The contractor shall refer all requests for, demands for production of or inquiries about, VA information and information systems to the VA CO for response.

- Notwithstanding the provision above, the contractor shall not release VA records protected by Title 38 U.S.C. § 5705, Confidentiality of medical quality- assurance records and/or Title 38 U.S.C. § 7332, Confidentiality of certain

medical records pertaining to drug addiction, sickle cell anemia, alcoholism or alcohol abuse or infection with Human Immunodeficiency Virus (HIV). If the contractor is in receipt of a court order or other requests for the above- mentioned information, the contractor shall immediately refer such court order or other requests to the VA CO for response.

- Information made available to the contractor by VA for the performance or administration of this contract or information developed by the contractor in performance or administration of the contract will be protected and secured in accordance with VA Directive 6500 and Identity and Access Management (IAM) Security processes specified in the VA Information Security Knowledge Service.

- Any data destruction done on behalf of VA by a contractor shall be done in accordance with National Archives and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records and Information Management, VA Handbook 6300.1, Records Management Procedures, and applicable VA Records Control Schedules.

- The contractor shall provide its plan for destruction of all VA data in its possession according to VA Directive 6500 and NIST 800-88, Guidelines for Media Sanitization prior to termination or completion of this contract. If directed by the POC /CO, the contractor shall return all Federal Records to VA for disposition.

- Any media, such as paper, magnetic tape, magnetic disks, solid state devices or optical discs that is used to store, process, or access VA information that cannot be destroyed shall be returned to VA. The contractor shall hold the appropriate material until otherwise directed by the Contracting Officer’s Representative (POC) or CO. Items shall be returned securely via VA-approved methods. VA sensitive information must be transmitted utilizing VA-approved encryption tools which are validated under FIPS 140-2 (or its successor) and NIST 800-52. If mailed, the contractor shall send via a trackable method (USPS, UPS, FedEx, etc.) and immediately provide the POC /CO with the tracking information. Self-certification by the contractor that the data destruction requirements above have been met shall be sent to the POC /CO within 30 business days of termination of the contract.

- All electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) used to store, process or access VA information will not be returned to the contractor at the end of lease, loan, or trade-in. Exceptions to this paragraph will only be granted with the written approval of the VA CO.

3. ACCESS TO VA INFORMATION AND VA INFORMATION SYSTEMS.

This section applies when any person requires access to information made available to the contractor by VA for the performance or administration of this contract or information developed by the contractor in performance or administration of the contract.

- A contractor/subcontractor shall request logical (technical) or physical access to VA information and VA information systems for their employees and subcontractors only to the extent necessary to perform the services specified in the solicitation or contract. This includes indirect entities, both affiliate of contractor/subcontractor and agent of contractor/subcontractor.

- Contractors and subcontractors shall sign the VA Information Security Rule of Behavior (ROB) before access is provided to VA information and information systems (see Section 4, Training, below). The ROB contains the minimum user compliance requirements and does not supersede any policies of VA facilities or other agency components which provide higher levels of protection to VA’s information or information systems. Users who require privileged access shall complete the VA elevated privilege access request processes before privileged access is granted.

- All contractors and subcontractors working with VA information are subject to the same security investigative and clearance requirements as those of VA appointees or employees who have access to the same types of information. The level and process of background security investigations for contractors shall be in accordance with VA Directive and Handbook 0710, Personnel Suitability and Security Program. The Office of Human Resources and Administration/Operations, Security and Preparedness (HRA/OSP) is responsible for these policies and procedures. Contract personnel who require access to classified information or information systems shall have an appropriate security clearance. Verification of a Security Clearance shall be processed through the Special Security Officer located in HRA/OSP. Contractors shall conform to all requirements stated in the National Industrial Security Program Operating Manual (NISPOM).

- All contractors and subcontractors shall comply with conditions specified in VAAR 852.204-71(d); Contractor operations required to be in United States. All contractors and subcontractors working with VA information must be permanently located within a jurisdiction subject to the law of the United States or its Territories to the maximum extent feasible. If services are proposed to be performed abroad the contractor must state where all non-U.S. services are provided. The contractor shall deliver to VA a detailed plan specifically addressing communications, personnel control, data protection and potential legal issues. The plan shall be approved by the POC /CO in writing prior to access being granted.

- The contractor shall notify the POC/CO in writing immediately (no later than 24 hours) after personnel separation or occurrence of other causes. Causes may include the following:

  - Contractor/subcontractor personnel no longer has a need for access to VA information or VA information systems.

- Contractor/subcontractor personnel are terminated, suspended, or otherwise has their work on a VA project discontinued for any reason.

- Contractor believes their own personnel or subcontractor personnel may pose a threat to their company’s working environment or to any company- owned property. This includes contractor-owned assets, buildings, confidential data, customers, employees, networks, systems, trade secrets and/or VA data.

- Any previously undisclosed changes to contractor/subcontractor background history are brought to light, including but not limited to changes to background investigation or employee record.

- Contractor/subcontractor personnel have their authorization to work in the United States revoked.

- Agreement by which contractor provides products and services to VA has either been fulfilled or terminated, such that VA can cut off electronic and/or physical access for contractor personnel.

- In such cases of contract fulfillment, termination, or other causes; the contractor shall take the necessary measures to immediately revoke access to VA network, property, information, and information systems (logical and physical) by contractor/subcontractor personnel. These measures include (but are not limited to): removing and then securing Personal Identity Verification (PIV) badges and PIV – Interoperable (PIV-I) access badges, VA-issued photo badges, credentials for VA facilities and devices, VA-issued laptops, and authentication tokens. Contractors shall notify the appropriate VA POC /CO immediately to initiate access removal.

- Contractors/subcontractors who no longer require VA accesses will return VA- issued property to VA. This property includes (but is not limited to): documents, electronic equipment, keys, and parking passes. PIV and PIV-I access badges shall be returned to the nearest VA PIV Badge Issuance Office. Once they have had access to VA information, information systems, networks and VA property in their possessions removed, contractors shall notify the appropriate VA POC /CO.

- TRAINING. This entire section applies to all acquisitions which include section 3.

  - All contractors and subcontractors requiring access to VA information and VA information systems shall successfully complete the following before being granted access to VA information and its systems:

    - VA Privacy and Information Security Awareness and Rules of Behavior course (Talent Management System (TMS) #10176) initially and annually thereafter.

- Sign and acknowledge (electronically through TMS #10176) understanding of and responsibilities for compliance with the Organizational Rules of Behavior, relating to access to VA information and information systems initially and annually thereafter; and

- Successfully complete any additional cyber security or privacy training, as required for VA personnel with equivalent information system or information access [to be defined by the VA program official and provided to the VA CO for inclusion in the solicitation document – i.e., any role- based information security training].

- The contractor shall provide to the POC /CO a copy of the training certificates and certification of signing the Organizational Rules of Behavior for each applicable employee within five days of the initiation of the contract and annually thereafter, as required.

- Failure to complete the mandatory annual training is grounds for suspension or termination of all physical or electronic access privileges and removal from work on the contract until such time as the required training is complete.

- SECURITY INCIDENT INVESTIGATION. This entire section applies to all acquisitions requiring any Information Security and Privacy language.

  - The contractor, subcontractor, their employees, or business associates shall immediately (within one hour) report suspected security / privacy incidents to the VA OIT’s Enterprise Service Desk (ESD) by calling (855) 673-4357 (TTY: 711). The ESD is OIT’s 24/7/365 single point of contact for IT-related issues. After reporting to the ESD, the contractor, subcontractor, their employees, or business associates shall, within one hour, provide the POC/CO the incident number received from the ESD.

  - To the extent known by the contractor/subcontractor, the contractor/ subcontractor's notice to VA shall identify the information involved and the circumstances surrounding the incident, including the following:

    - The date and time (or approximation of) the Security Incident occurred.

    - The names of individuals involved (when applicable).

    - The physical and logical (if applicable) location of the incident.

    - Why the Security Incident took place (i.e., catalyst for the failure).

    - The amount of data belonging to VA believed to have been compromised.

    - The remediation measures the contractor is taking to ensure no future incidents of a similar nature.

- After the contractor has provided the initial detailed incident summary to VA, they will continue to provide written updates on any new and relevant circumstances or facts they discover. The contractor, subcontractor, and their employes shall fully cooperate with VA or third-party entity performing an independent risk analysis on behalf of VA. Failure to cooperate may be deemed a material breach and grounds for contract termination.

- VA IT contractors shall follow VA Handbook 6500, Risk Management Framework for VA Information Systems VA Information Security Program, and VA Information Security Knowledge Service guidance for implementing an Incident Response Plan or integrating with an existing VA implementation.

- In instances of theft or break-in or other criminal activity, the contractor/subcontractor must concurrently report the incident to the appropriate law enforcement entity (or entities) of jurisdiction, including the VA OIG, and the VA Office of Security and Law Enforcement. The contractor, its employees, and its subcontractors and their employees shall cooperate with VA and any law enforcement authority responsible for the investigation and prosecution of any possible criminal law violation(s) associated with any incident. The contractor/subcontractor shall cooperate with VA in any civil litigation to recover VA information, obtain monetary or other compensation from a third party for damages arising from any incident, or obtain injunctive relief against any third party arising from, or related to, the incident.

- The contractor shall comply with VA Handbook 6500.2, Management of Breaches Involving Sensitive Personal Information, which establishes the breach management policies and assigns responsibilities for the oversight, management and reporting procedures associated with managing of breaches.

- With respect to unsecured Protected Health Information (PHI), the contractor is deemed to have discovered a data breach when the contractor knew or should have known of breach of such information. When a business associate is part of VHA contract, notification to the covered entity (VHA) shall be made in accordance with the executed BAA.

- If the contractor or any of its agents fails to protect VA sensitive personal information or otherwise engages in conduct which results in a data breach involving any VA sensitive personal information the contractor/subcontractor processes or maintains under the contract; the contractor shall pay liquidated damages to the VA as set forth in clause 852.211-76, Liquidated Damages— Reimbursement for Data Breach Costs.

6. INFORMATION SYSTEM DESIGN AND DEVELOPMENT. This entire section

applies to information systems, systems, major applications, minor applications, enclaves, and platform information technologies (to include the subcomponents of each) designed or developed for or on behalf of VA by any non-VA entity.

- Information systems designed or developed on behalf of VA at non-VA facilities shall comply with all applicable Federal law, regulations, and VA policies. This includes standards for the protection of electronic Protected Health Information (PHI), outlined in 45 C.F.R. Part 164, Subpart C and information and system security categorization level designations in accordance with FIPS 199, Standards for Security Categorization of Federal Information and Information Systems and FIPS 200, Minimum Security Requirements for Federal Information Systems. Baseline security controls shall be implemented commensurate with the FIPS 199 system security categorization (reference VA Handbook 6500 and VA Trusted Internet Connections (TIC) Architecture).

- Contracted new developments require creation, testing, evaluation, and authorization in compliance with VA Assessment and Authorization (A&A) processes in VA Handbook 6500 and VA Information Security Knowledge Service to obtain an Authority to Operate (ATO). VA Directive 6517, Risk Management Framework for Cloud Computing Services, provides the security and privacy requirements for cloud environments.

- VA IT contractors, subcontractors and third-party service providers shall address and/or integrate applicable VA Handbook 6500, VA Handbook 6517, Risk Management Framework for Cloud Computing Services and Information Security Knowledge Service specifications in delivered IT systems/solutions, products and/or services. If systems/solutions, products and/or services do not directly match VA security requirements, the contractor shall work though the POC/CO to identify the VA organization responsible for governance or resolution. Contractors shall comply with RFO 39.1, specifically the prohibitions referenced.

- The contractor (including producers and resellers) shall comply with Office of Management and Budget (OMB) M-22-18 and M-23-16 when using third-party software on VA information systems or otherwise affecting the VA information. This includes new software purchases and software renewals for software developed or modified by major version change after the issuance date of M- 22-18 (September 14, 2022). The term “software” includes firmware, operating systems, applications and application services (e.g., cloud-based software), as well as products containing software. The contractor shall provide a self- attestation that secure software development practices are utilized as outlined by Executive Order (EO)14028 and NIST Guidance. A third-party assessment provided by either a certified Federal Risk and Authorization Management Program (FedRAMP) Third Party Assessor Organization (3PAO) or one approved by the agency will be acceptable in lieu of a software producer's self- attestation.

- The contractor shall ensure all delivered applications, systems and information systems are compliant with Homeland Security Presidential Directive (HSPD) 12 and VA Identity and Access management (IAM) enterprise identity management requirements as set forth in OMB M-19-17, M-05-24, FIPS 201-3,

Personal Identity Verification (PIV) of Federal Employees and Contractors (or its successor), M-21-31 and supporting NIST guidance. This applies to Commercial Off-The-Shelf (COTS) product(s) that the contractor did not develop, all software configurations and all customizations.

- The contractor shall ensure all contractor delivered applications and systems provide user authentication services compliant with VA Handbook 6500, VA Information Security Knowledge Service, IAM enterprise requirements and NIST 800-63, Digital Identity Guidelines, for direct, assertion-based authentication and/or trust-based authentication, as determined by the design and integration patterns. Direct authentication at a minimum must include Public Key Infrastructure (PKI) based authentication supportive of PIV and/or Common Access Card (CAC), as determined by the business need and compliance with VA Information Security Knowledge Service specifications.

- The contractor shall use VA authorized technical security baseline configurations and certify to the POC that applications are fully functional and operate correctly as intended on systems in compliance with VA baselines prior to acceptance or connection into an authorized VA computing environment. If the Defense Information Systems Agency (DISA) has created a Security Technical Implementation Guide (STIG) for the technology, the contractor may configure to comply with that STIG. If VA determines a new or updated VA configuration baseline needs to be created, the contractor shall provide required technical support to develop the configuration settings.

- The standard installation, operation, maintenance, updating and patching of software shall not alter the configuration settings from VA approved baseline configuration. Software developed for VA must be compatible with VA enterprise installer services and install to the default “program files” directory with silently install and uninstall. The contractor shall perform testing of all updates and patching prior to implementation on VA systems.

- Applications designed for normal end users will run in the standard user context without elevated system administration privileges.

- The contractor-delivered solutions shall reside on VA approved operating systems. Exceptions to this will only be granted with the written approval of the POC/CO.

- The contractor shall design, develop, and implement security and privacy controls in accordance with the provisions of VA security system development life cycle outlined in NIST 800-37, Risk Management Framework for Information Systems and Organizations: A System Life Cycle Approach for Security and Privacy, VA Directive and Handbook 6500, and VA Handbook 6517.

- The Contractor shall comply with the Privacy Act of1974 (the Act), RFO 52.224-2 Privacy Act, and VA rules and regulations issued under the Act in the design, development, or operation of any system of records on individuals to accomplish a VA function.

- The contractor shall ensure the security of all procured or developed information systems, systems, major applications, minor applications, enclaves and platform information technologies, including their subcomponents (hereinafter referred to as “Information Systems”) throughout the life of this contract and any extension, warranty, or maintenance periods. This includes security configurations, workarounds, patches, hotfixes, upgrades, replacements and any physical components which may be necessary to remediate all security vulnerabilities published or known to the contractor anywhere in the information systems (including systems, operating systems, products, hardware, software, applications and firmware). The contractor shall ensure security fixes do not negatively impact the Information Systems.

- When the contractor is responsible for operations or maintenance of the systems, the contractor shall apply the security fixes within the timeframe specified by the associated controls on the VA Information Security Knowledge Service. When security fixes involve installing third party patches (such as Microsoft OS patches or Adobe Acrobat), the contractor shall provide written notice to the VA POC/CO that the patch has been validated as to not affecting the Systems within 10 business days.

- INFORMATION SYSTEM HOSTING, OPERATION, MAINTENANCE OR USE.

This entire section applies to information systems, systems, major applications, minor applications, enclaves, and platform information technologies (cloud and non- cloud) hosted, operated, maintained, or used on behalf of VA at non-VA facilities.

  - The contractor shall comply with all Federal laws, regulations, and VA policies for Information systems (cloud and non-cloud) that are hosted, operated, maintained, or used on behalf of VA at non-VA facilities. Security controls for collecting, processing, transmitting, and storing of VA sensitive information, must be in place. The controls will be tested by VA or a VA sanctioned 3PAO and approved by VA prior to hosting, operation, maintenance or use of the information system or systems by or on behalf of VA. This includes conducting compliance risk assessments, security architecture analysis, routine vulnerability scanning, system patching, change management procedures and the completion of an acceptable contingency plan for each system. The contractor’s security control procedures shall be the same as procedures used to secure VA-operated information systems.

  - Outsourcing (contractor facility, equipment, or staff) of systems or network operations, telecommunications services or other managed services require Assessment and Authorization (A&A) of the contractor’s systems in accordance with VA Handbook 6500 as specified in VA Information Security Knowledge

Service. Major changes to the A&A package may require reviewing and updating all the documentation associated with the change. The contractor’s cloud computing systems shall comply with FedRAMP and VA Directive 6517 requirements.

- The contractor shall return all electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) on non-VA leased or non-VA owned IT equipment used to store, process or access VA information to VA in accordance with A&A package requirements. This applies when the contract is terminated or completed and prior to disposal of media. The contractor shall provide its plan for destruction of all VA data in its possession according to VA Information Security Knowledge Service requirements and NIST 800-88. The contractor shall send a self-certification that the data destruction requirements above have been met to the POC/CO within 30 business days of termination of the contract.

- All external internet connections to VA network involving VA information must be in accordance with VA Trusted Internet Connection (TIC) Reference Architecture and VA Directive and Handbook 6513, Secure External Connections and reviewed and approved by VA prior to implementation. Government-owned contractor-operated systems, third party or business partner networks require a Memorandum of Understanding (MOU) and Interconnection Security Agreements (ISA).

- Contractor procedures shall be subject to periodic, announced, or unannounced assessments by VA officials, the OIG or a 3PAO. The physical security aspects associated with contractor activities are also subject to such assessments. The contractor shall report, in writing, any deficiencies noted during the above assessment to the VA POC/CO. The contractor shall use VA’s defined processes to document planned remedial actions that address identified deficiencies in information security policies, procedures, and practices. The contractor shall correct security deficiencies within the timeframes specified in the VA Information Security Knowledge Service.

- All major information system changes which occur in the production environment shall be reviewed by the VA to determine the impact on privacy and security of the system. Based on the review results, updates to the Authority to Operate (ATO) documentation and parameters may be required to remain in compliance with VA Handbook 6500 and VA Information Security Knowledge Service requirements.

- The contractor shall conduct an annual privacy and security self-assessment on all information systems and outsourced services as required. Copies of the assessment shall be provided to the POC/CO. The VA/Government reserves the right to conduct assessment using government personnel or a third-party if deemed necessary. The contractor shall correct or mitigate any weaknesses discovered during the assessment.

- VA prohibits the installation and use of personally owned or contractor-owned equipment or software on VA information systems. If non-VA owned equipment must be used to fulfill the requirements of a contract, it must be stated in the service agreement, SOW, PWS, PD or contract. All security controls required for government furnished equipment must be utilized in VA approved Other Equipment (OE). Configuration changes to the contractor OE, must be funded by the owner of the equipment. All remote systems must use a VA-approved antivirus software and a personal (host-based or enclave based) firewall with a VA-approved configuration. The contractor shall ensure software on OE is kept current with all critical updates and patches. Owners of approved OE are responsible for providing and maintaining the anti-virus software and the firewall on the non-VA owned OE. Approved contractor OE will be subject to technical inspection at any time.

- The contractor shall notify the POC/CO within one hour of disclosure or successful exploits of any vulnerability which can compromise the confidentiality, integrity, or availability of the information systems. The system or effected component(s) need(s) to be isolated from the network. A forensic analysis needs to be conducted jointly with VA. Such issues will be remediated as quickly as practicable, but in no event longer than the timeframe specified by VA Information Security Knowledge Service. If sensitive personal information is compromised reference VA Handbook 6500.2 and Section 5, Security Incident Investigation.

- For cases wherein the contractor discovers material defects or vulnerabilities impacting products and services they provide to VA, the contractor shall develop and implement policies and procedures for disclosure to VA, as well as remediation. The contractor shall, within 30 business days of discovery, document a summary of these vulnerabilities or defects. The documentation will include a description of the potential impact of each vulnerability and material defect, compensating security controls, mitigations, recommended corrective actions, root cause analysis and/or workarounds (i.e., monitoring). Should there exist any backdoors in the products or services they provide to VA (referring to methods for bypassing computer authentication), the contractor shall provide the VA CO/CO written assurance they have permanently remediated these backdoors.

- All other vulnerabilities, including those discovered through routine scans or other assessments, will be remediated based on risk, in accordance with the remediation timelines specified by the VA Information Security Knowledge Service and/or the applicable timeframe mandated by Cybersecurity & Infrastructure Security Agency (CISA) Binding Operational Directive (BOD) 22- 01 and BOD 19-02 for Internet-accessible systems. Exceptions to this paragraph will only be granted with the approval of the POC/CO.

# 8. SECURITY AND PRIVACY CONTROLS COMPLIANCE TESTING, ASSESSMENT

AND AUDITING. This entire section applies whenever section 6 or 7 is included.

- Should VA request it, the contractor shall provide a copy of their (corporation’s, sole proprietorship’s, partnership’s, limited liability company (LLC), or other business structure entity’s) policies, procedures, evidence and independent report summaries related to specified cybersecurity frameworks (International Organization for Standardization (ISO), NIST Cybersecurity Framework (CSF), etc.). VA or its third-party/partner designee (if applicable) are further entitled to perform their own audits and security/penetration tests of the contractor’s IT or systems and controls, to ascertain whether the contractor is complying with the information security, network or system requirements mandated in the agreement between VA and the contractor.

- Any audits or tests of the contractor or third-party designees/partner VA elects to carry out will commence within 30 business days of VA notification. Such audits, tests and assessments may include the following: (a): security/penetration tests which both sides agree will not unduly impact contractor operations; (b): interviews with pertinent stakeholders and practitioners; (c): document review; and (d): technical inspections of networks and systems the contractor uses to destroy, maintain, receive, retain, or use VA information.

- As part of these audits, tests and assessments, the contractor shall provide all information requested by VA. This information includes, but is not limited to, the following: equipment lists, network or infrastructure diagrams, relevant policy documents, system logs or details on information systems accessing, transporting, or processing VA data.

- The contractor and at its own expense, shall comply with any recommendations resulting from VA audits, inspections and tests. VA further retains the right to view any related security reports the contractor has generated as part of its own security assessment. The contractor shall also notify VA of the existence of any such security reports or other related assessments, upon completion and validation.

- VA appointed auditors or other government agency partners may be granted access to such documentation on a need-to-know basis and coordinated through the POC/CO. The contractor shall comply with recommendations which result from these regulatory assessments on the part of VA regulators and associated government agency partners.

- PRODUCT INTEGRITY, AUTHENTICITY, PROVENANCE, ANTI-COUNTERFEIT

AND ANTI-TAMPERING. This entire section applies when the acquisition involves any product (application, hardware, or software) or when section 6 or 7 is included.

  - The contractor shall comply with Code of Federal Regulations (CFR) Title 15 Part 7, “Securing the Information and Communications Technology and Services (ICTS) Supply Chain”, which prohibits ICTS Transactions from foreign adversaries. ICTS Transactions are defined as any acquisition, importation, transfer, installation, dealing in or use of any information and communications technology or service, including ongoing activities, such as managed services, data transmission, software updates, repairs or the platforming or data hosting of applications for consumer download.

- When contracting terms require the contractor to procure equipment, the contractor shall purchase or acquire the equipment from an Original Equipment Manufacturer (OEM) or an authorized reseller of the OEM. The contractor shall attest that equipment procured from an OEM or authorized reseller or distributor are authentic. If procurement is unavailable from an OEM or authorized reseller, the contractor shall submit in writing, details of the circumstances prohibiting this from happening and procure a product waiver from the VA POC/CO.

- All contractors shall establish, implement, and provide documentation for risk management practices for supply chain delivery of hardware, software (to include patches) and firmware provided under this agreement. Documentation will include chain of custody practices, inventory management program, information protection practices, integrity management program for sub-supplier provided components, and replacement parts requests. The contractor shall make spare parts available. All contractor(s) shall specify how digital delivery for procured products, including patches, will be validated and monitored to ensure consistent delivery. The contractor shall apply encryption technology to protect procured products throughout the delivery process.

- If a contractor provides software or patches to VA, the contractor shall publish or provide a hash conforming to the FIPS Security Requirements for Cryptographic Modules (FIPS 140-2 or successor).

- The contractor shall provide a software bill of materials (SBOM) for procured (to include licensed products) and consist of a list of components and associated metadata which make up the product. SBOMs must be generated in one of the data formats defined in the National Telecommunications and Information Administration (NTIA) report “The Minimum Elements for a Software Bill of Materials (SBOM).”

- Contractors shall use or arrange for the use of trusted channels to ship procured products, such as U.S. registered mail and/or tamper-evident packaging for physical deliveries.

- Throughout the delivery process, the contractor shall demonstrate a capability for detecting unauthorized access (tampering).

- The contractor shall demonstrate chain-of-custody documentation for procured products and require tamper-evident packaging for the delivery of this hardware.

- VIRUSES, FIRMWARE AND MALWARE. This entire section applies when the acquisition involves any product (application, hardware, or software) or when section 6 or 7 is included.

  - The contractor shall execute due diligence to ensure all provided software and patches, including third-party patches, are free of viruses and/or malware before releasing them to or installing them on VA information systems.

  - The contractor warrants it has no knowledge of and did not insert, any malicious virus and/or malware code into any software or patches provided to VA which could potentially harm or disrupt VA information systems. The contractor shall use due diligence, if supplying third-party software or patches, to ensure the third-party has not inserted any malicious code and/or virus which could damage or disrupt VA information systems.

  - The contractor shall provide or arrange for the provision of technical justification as to why any “false positive” hit has taken place to ensure their code’s supply chain has not been compromised. Justification may be required, but is not limited to, when install files, scripts, firmware, or other contractor-delivered software solutions (including third-party install files, scripts, firmware, or other software) are flagged as malicious, infected, or suspicious by an anti-virus vendor.

  - The contractor shall not upload (intentionally or negligently) any virus, worm, malware or any harmful or malicious content, component and/or corrupted data/source code (hereinafter “virus or other malware”) onto VA computer and information systems and/or networks. If introduced (and this clause is violated), upon written request from the VA CO, the contractor shall:

    - Take all necessary action to correct the incident, to include any and all assistance to VA to eliminate the virus or other malware throughout VA’s information networks, computer systems and information systems; and

    - Use commercially reasonable efforts to restore operational efficiency and remediate damages due to data loss or data integrity damage, if the virus or other malware causes a loss of operational efficiency, data loss, or damage to data integrity.

- CRYPTOGRAPHIC REQUIREMENT. This entire section applies whenever the acquisition includes section 6 or 7 is included.

  - The contractor shall document how the cryptographic system supporting the contractor’s products and/or services protect the confidentiality, data integrity, authentication and non-repudiation of devices and data flows in the underlying system.

- The contractor shall use only approved cryptographic methods as defined in FIPS 140-2 (or its successor) and NIST 800-52 standards when enabling encryption on its products.

- The contractor shall provide or arrange for the provision of an automated remote key-establishment method which protects the confidentiality and integrity of the cryptographic keys.

- The contractor shall ensure emergency re-keying of all devices can be remotely performed within 30 business days.

- The contractor shall provide or arrange for the provision of a method for updating cryptographic primitives or algorithms.

- PATCHING GOVERNANCE. This entire section applies whenever the acquisition includes section 7 is included

  - The contractor shall provide documentation detailing the patch management, vulnerability management, mitigation and update processes (to include third- party) prior to the connection of electronic devices, assets or equipment to VA’s assets. This documentation will include information regarding the follow:

    - The resources and technical capabilities to sustain the program or process (e.g., how the integrity of a patch is validated by VA); and

    - The approach and capability to remediate newly reported zero-day vulnerabilities for contractor products.

  - The contractor shall verify and provide documentation all procured products (including third-party applications, hardware, software, operating systems, and firmware) have appropriate updates and patches installed prior to delivery to VA.

  - The contractor shall provide or arrange the provision of appropriate software and firmware updates to remediate newly discovered vulnerabilities or weaknesses for their products and services within 30 days of discovery. Updates to remediate critical or emergent vulnerabilities will be provided within seven business days of discovery. If updates cannot be made available by contractor within these time periods, the contractor shall submit mitigations, methods of exploit detection and/or workarounds to the POC/CO prior to the above deadlines.

  - The contractor shall provide or arrange for the provision of appropriate hardware, software and/or firmware updates, when those products, including open-source software, are provided to the VA, to remediate newly discovered vulnerabilities or weaknesses. Remediations of products or services provided to the VA’s system environment must be provided within 30 business days of availability from the original supplier and/or patching source. Updates to remediate critical vulnerabilities applicable to the Contractor’s use of the third- party product in its system environment will be provided within seven business days of availability from the original supplier and/or patching source. If applicable third-party updates cannot be integrated, tested and made available by Contractor within these time periods, mitigations and/or workarounds will be provided to the POC/CO before the above deadlines.

- SPECIALIZED DEVICES/SYSTEMS (MEDICAL DEVICES, SPECIAL PURPOSE

SYSTEMS, RESEARCH SCIENTIFIC COMPUTING). This entire section applies when the acquisition includes one or more Medical Device, Special Purpose System or Research Scientific Computing Device. If appropriate, ensure selected clauses from section 6 or 7 and 8 through 12 are included.

  - Contractor supplies/delivered Medical Devices, Special Purpose Systems- Operational Technology (SPS-OT) and Research Scientific Computing Devices shall comply with all applicable Federal law, regulations, and VA policies. New developments require creation, testing, evaluation, and authorization in compliance with processes specified on the Specialized Device Cybersecurity Department Enterprise Risk Management (SDCD-ERM) Portal, VA Directive 6550, Pre-Procurement Assessment and Implementation of Medical Devices/Systems, VA Handbook 6500, and the VA Information Security Knowledge Service. Deviations from Federal law, regulations, and VA Policy are identified and documented as part of VA Directive 6550 and/or the VA Enterprise Risk Analysis (ERA) processes for Specialized Devices/Systems processes.

  - All contractors and third-party service providers shall address and/or integrate applicable VA Handbook 6500 and Information Security Knowledge Service specifications in delivered IT systems/solutions, products and/or services. If systems/solutions, products and/or services do not directly match VA security requirements, the contractor shall work though the POC/CO for governance or resolution.

  - The contractor shall certify to the POC/CO that devices/systems that have completed the VA Enterprise Risk Analysis (ERA) process for Specialized Devices/Systems are fully functional and operate correctly as intended. Devices/systems must follow the VA ERA authorized configuration prior to acquisition and connection to the VA computing environment. If VA determines a new VA ERA needs to be created, the contractor shall provide required technical support to develop the configuration settings. Major changes to a previously approved device/system will require a new ERA.

  - The contractor shall comply with all practices documented by the Food Drug and Administration (FDA) Premarket Submission for Management of Cybersecurity in Medical Devices and Postmarket Management of Cybersecurity in Medical Devices.

- The contractor shall design devices capable of accepting all applicable security patches with or without the support of the contractor personnel. If patching can only be completed by the contractor, the contractor shall commit the resources needed to patch all applicable devices at all VA locations. If unique patching instructions or packaging is needed, the contractor shall provide the necessary information in conjunction with the validation/testing of the patch. The contractor shall apply security patches within 30 business days of the patch release and have a formal tracking process for any security patches not implemented to include explanation when a device cannot be patched.

- The contractor shall provide devices able to install and maintain VA-approved antivirus capabilities with the capability to quarantine files and be updated as needed in response to incidents. Alternatively, a VA-approved whitelisting application may be used when the contractor cannot install an anti-virus / anti- malware application.

- The contractor shall verify and document all software embedded within the device does not contain any known viruses or malware before delivery to or installation at a VA location.

- Devices and other equipment or systems containing media (hard drives, optical disks, solid state, and storage via chips/firmware) with VA sensitive

information will be returned to the contractor with media removed. When the contract requires return of equipment, the options available to the contractor are the following:

  - The contractor shall accept the system without the drive, firmware and solid state.

  - VA’s initial device purchase includes a spare drive or other replacement media which must be installed in place of the original drive at time of turn- in; or

  - Due to the highly specialized and sometimes proprietary hardware and software associated with the device, if it is not possible for VA to retain the hard drive, firmware, and solid state, then:

    - The equipment contractor shall have an existing BAA if the device being traded in has sensitive information stored on it and hard drive(s) from the system are being returned physically intact.

    - Any fixed hard drive, Complementary Metal-Oxide-Semiconductor (CMOS), Programmable Read-Only Memory (PROM), solid state and firmware on the device must be non-destructively sanitized to the greatest extent possible without negatively impacting system operation. Selective clearing down to patient data folder level is recommended using VA approved and validated overwriting technologies/methods/tools. Applicable media sanitization specifications need to be pre-approved and described in the solicitation, contract, or order.

- DATA CENTER PROVISIONS. This entire section applies whenever the acquisition requires an interconnection to/from the VA network to/from a non-VA location.

  - The contractor shall ensure the VA network is accessed by in accordance with VA Directive 6500 and IAM security processes specified in the VA Information Security Knowledge Service.

  - The contractor shall ensure network infrastructure and data availability in accordance with VA information system business continuity procedures specified in the VA Information Security Knowledge Service.

  - The contractor shall ensure any connections to the internet or other external networks for information systems occur through managed interfaces utilizing VA approved boundary protection devices (e.g., internet proxies, gateways, routers, firewalls, guards or encrypted tunnels).

  - The contractor shall encrypt all traffic across the segment of the Wide Area Network (WAN) it manages and no unencrypted Out of Band (OOB) Internet Protocol (IP) traffic will traverse the network.

  - The contractor shall ensure tunnel endpoints are routable addresses at each VA operating site.

  - The contractor shall secure access from Local Area Networks (LANs) at co- located sites in accordance with VA TIC Reference Architecture, VA Directive and Handbook 6513, and MOU/ISA process specified in the VA Information Security Knowledge Service.