| SUBJECT* | Tower and Scope Lease, Onsite Endoscopy Specialist, PM and Repair |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 23667 |
| --- | --- |
| SOLICITATION NUMBER* | 36C24626Q0530 |
| RESPONSE DATE/TIME/ZONE | 05-03-2026 6:00pm EASTERN TIME, NEW YORK, USA |
| ARCHIVE | 7 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | W065 |
| NAICS CODE* | 339113 |
| CONTRACTING OFFICE ADDRESS | Hampton VAMC Department of Veterans Affairs Network Contracting Office 6 100 Emancipation Drive Hampton VA 23667 |
| POINT OF CONTACT* | Contracting Specialist Stephen Nassan stephen.nassan@va.gov 7573152910 |

PLACE OF PERFORMANCE

| ADDRESS | Richmond VAMC |
| --- | --- |
|  | 1201 Broad Rock Blvd |
|  | Richmond VA |
| POSTAL CODE | 23249 |
| COUNTRY | USA |

ADDITIONAL INFORMATION

| AGENCY’S URL | https://www.va.gov |
| --- | --- |
| URL DESCRIPTION | https://www.va.gov |
| AGENCY CONTACT’S EMAIL ADDRESS | stephen.nassan@va.gov |
| EMAIL DESCRIPTION | stephen.nassan@va.gov |

DESCRIPTION

This “Sources Sought Notice” is for the Richmond VAMC to determine sources for Tower and Scope Lease, Onsite Endoscopy Specialist, PM and Repair listed as the following:

9 or more units of - 4K/ICG/3D compatible Video Systems

2 or more units of - Video Mediastinascopes

2 or more units of - 4K/3D Exscope System

1 or more units of – Service plan for the 13 units

2 units of – Onsite specialist for the 13 or more units

Salient Characteristics are as follows:

System Functions & Performance

- Configuration & Integration:

  - Supports multiple configuration options for various operating room (OR) layouts.

  - Modular and expandable components for future upgrades.

  - Integrates with VA IT systems (such as VistA or Cerner) and platforms like StreamConnect.

  - Interfaces with a broad range of endoscopic and surgical devices.

  - Scopes, cameras, and instruments are interchangeable and cross-compatible with tower and software.

- Stability & Ergonomics:

  - Tower arms/supports stably hold up to 66 lbs.

  - Multiple monitor/light mounting points.

  - User interfaces/control panels accessible from standing position; operable with gloved hands.

  - Ergonomic design to reduce staff fatigue.

    - Lightweight cameras (≤ 350 grams).

- Cable Management:

  - Integrated system minimizing visible cables to no more than two (≤ 2).

  - Cables accessible for adjustments/repair within 5 minutes.

- Hygiene & Safety:

  - Non-porous surfaces, cleanable/disinfectable per CDC/VA standards in under 5 minutes.

  - Built-in locks or safety features to prevent accidental movement/dislodgement.

  - Scopes and surgical instruments are fully compatible with autoclave sterilization (per VA Directive 1116).

- Imaging & Compatibility:

  - Cameras support HD and 4K video output (≥3840×2160 resolution).

  - Supports 3D/4K exoscopes for neurosurgery/ENT procedures.

  - Camera Control Unit (CCU) backward/forward compatible with all VA camera heads/scopes.

  - Supports rigid, flexible, exoscopic, and disposable scopes for airway, ENT, urology, and more.

  - Imaging enhancement features:

    - Dynamic brightness adjustment.

    - Red contrast enhancement.

    - Near-infrared (NIR) imaging with ICG fluorescence.

    - Narrow-band imaging (NBI).

    - True-color reproduction.

    - Wide field of view.

  - Integrated image/video recording.

  - Secure, VA-compliant data transfer to VA IT platforms (encrypted wireless transfer if required).

- Durability:

  - Instruments/scopes/accessories compatible with repeated autoclave sterilization.

Service & Support

- On-site technical support (endoscopy specialist/technician) available during standard operating hours (minimally 7:00am–6:00pm).

- Telephone technical support available 24/7.

- On-site technician supports training, live troubleshooting, and rapid issue resolution.

- Replacement equipment (scopes, cameras, instruments) provided within 48 hours of notification.

- Initial and ongoing training for OR and clinical staff; in-service dates scheduled within 48 hours of request and delivered within seven days.

- Preventive and corrective maintenance performed by OEM-certified personnel per manufacturer’s recommendations.

- Repairs documented and reported to VA Biomedical Engineering, including service reports/test data.

- Functionally equivalent loaner system/component provided within 48 hours of system failure.

- Updates, patching, and technology refreshes per OEM specification, at VA Biomedical Engineering’s discretion.

- Compliance with VA infection control and patient privacy policies.

Physical & Dimensional Requirements

- Designed for mounting to ceiling booms to maximize floor space.

- Carts/towers fit through standard VA doorways (≤36” wide).

- Supports all listed scopes/instruments within manufacturer’s recommended weight/dimensions.

Clinical Compatibility & Special Features

- Includes video mediastinoscope for thoracic surgery.

- Supports 4K/3D exoscope system(s) for ENT/neurosurgery.

- Cross-functional use: towers, cameras, and scopes are interchangeable for Urology, Gastroenterology, General Surgery, ENT, Thoracic Surgery, and selected Neurosurgery.

Accessories & Future Upgrades

- Supports mounting/integration of additional VA-approved OR accessories and add-ons.

- Allows for future technology upgrades (hardware/software) during service agreement.

Compliance & General Provisions

- All work compliant with VA, OSHA, and facility safety regulations.

- Equipment compatible with VA security/privacy requirements (HIPAA, VA Handbook 6500).

- Service staff trained/certified on relevant equipment.

- All warranty/service agreement terms clearly outlined in proposal/submission.

ADDITIONAL INFORMATION:

This is a SOURCES SOUGHT announcement only. The purpose of this notice is to conduct market research and obtain information from qualified businesses. This is not a request for bids, quotes, or proposals. All responsible sources must submit a capability statement addressing their capabilities, capacity, place of manufacture, and location which shall be considered by the agency.

This notice shall not be construed as a commitment by the Government to ultimately award a contract, nor does it restrict the Government to a particular acquisition approach. All information submitted in response to this announcement is voluntary; the Government will not pay for information requested nor will it compensate any respondent for any cost incurred in developing information provided to the Government.

Interested parties submitting a response to this inquiry must include the following information:

- Capability statement

- Name and address of company

- NAICS Code

- SAM.gov UEI number

- Business size (large, small, VOSB, or SDVOSB)

- Place of manufacture (if the requirement involves supplies)

- Identify applicable socioeconomic categories and any pertinent information which demonstrates the firm’s ability to meet the above requirement

- An authorized distributor letter if the interested party is not the manufacturer

All responses shall be submitted no later than Sunday, 05/01/2026, by 06:00 PM EST.

Responses shall be emailed to: Stephen.Nassan@va.gov. When replying, please include “Sources Sought 36C24626Q0530” to ensure the email is reviewed in a timely manner.