PERFORMANCE WORK STATEMENT

Biomedical Engineering Technical Support Services

Sample Task Order - OEHRM

VHA HEALTHCARE TECHNOLOGY MANAGEMENT

NETWORKED MEDICAL DEVICE DATABASE 2.0 GUIDEBOOK

VERSION 1

XXXX XX,, 2018

Contents

PERFORMANCE WORK STATEMENT (PWS)	4

1.	Introduction	4

2.	Scope of Work	4

3.	Period of Performance	5

4.	Place of Performance	5

5.	Hours of Performance	5

6.	Travel	5

7.	Scheduling and Availability Reports	5

7.1 Kick-Off Meeting:	6

7.2 Monthly Status Reports:	6

8.	List of Equipment	6

9.	Services Overview	6

9.1.	Planned Maintenance (PM)	11

9.2.	Corrective Maintenance	12

9.3.	Professional Services	12

9.4.	Remote Support	15

9.5.	Test Equipment	15

9.6.	Parts and Service	15

10.	Work Order Process	15

11.	Physical Demands	16

12.	Work Environment	17

13.	Policies, Safety and Infection Control	17

14.	Contractor Key Personnel	17

14.1.	Basic Qualifications	17

14.2.	Role-Specific Qualifications	17

14.3.	Evidence	22

14.4.	Staffing Plan	22

14.5.	Acceptance	22

15.	Badging and Security	23

16.	Contractor Performance	23

17.	Information Security/Privacy for Inclusion into Contracts	24

17.1	General	24

17.2	Access to VA Information and VA Information Systems	25

17.3	VA Information Custodial Language	25

17.4	Information System Design and Development	27

17.5	Information System Hosting, Operation, Maintenance, Or Use	29

17.6	Security Incident Investigation	31

17.7	Liquidated Damages for Data Breach	31

17.8	Security Controls Compliance Testing	33

17.9	Training	33

18.	Records Management Language for Contracts	33

Attachment A: List of Facilities	36

Attachment B: List of Types of Medical Equipment	36

Attachment E: Pricing Spreadsheet	36

## PERFORMANCE WORK STATEMENT (PWS)

## Introduction

Veterans Health Administration (VHA) Office of Healthcare Technology Management (HTM) develops and guides the comprehensive management of healthcare technologies to ensure safe, available, and innovative medical technology used to deliver exceptional health care for Veterans.  HTM’s vision is to be an efficient, innovative and customer-oriented Center of Excellence that continually enhances VHA’s ability to deliver world class health care for our nation’s Veterans through exceptional management of healthcare technology as well as position VHA as the national leader in healthcare technology management. HTM/Biomedical Engineering at the medical facilities and VISNs direct the medical equipment management program and support clinical care through providing preventive and corrective maintenance, healthcare consulting, and professional services for the technology assessment, implementation and sustainment of medical equipment and clinical systems.

## Scope of Work

The purpose of this contract is to provide technical and professional support to Biomedical Engineering departments at Department of Veterans Affairs (VA) medical centers, outpatient clinics, VISN offices, and the HTM Program Office to accomplish corrective and planned maintenance, professional Biomedical Engineering services, and project management work.  Contract resources may be required to accomplish routine operational tasks (e.g. maintenance, patient safety recall investigation), as well as unexpected or time-limited tasks (e.g. COVID-19 pandemic support, operational readiness services for Office of Electronic Health Record Modernization (OEHRM)/Cerner deployment).  There are many competing priorities for Biomedical Engineering staff in today’s environment, and the support services shall provide interim Biomedical Engineering services either onsite or remotely while full-time positions are recruited and hired to fill permanent roles in Biomedical Engineering departments, as needed.  When Contractor support services are required by a VA medical facility, a local VISN or facility VA Contracting Officer Representative (L-COR) will be designated to monitor and provide oversight to the Contractors at that facility.

Successful offeror(s) may have to execute a Business Associate Agreement (BAA). Once a contract requiring a BAA is awarded, please contact the VHA Privacy Office BAA Issues team at VHABAAIssues@va.gov to assist in implementing the appropriate BAA.

These services are non-personal. Contractor personnel shall serve in a support role and do not make final decisions, approvals, or determinations on behalf of the Government. Contractor personnel are not subject to the supervision and control of a Government officer or employee in accordance with (IAW) FAR 37.201-2. Rather, Contractor personnel shall perform their duties IAW specific tasks outlined in the PWS. Supervisory functions such as hiring, directing, counseling, and firing of Contractor personnel are not performed by the Government. Contractor personnel shall at no time allow an employer-employee relationship to develop with VA organizations or their staff. VA organizations will refrain from any activities that create the appearance of such a relationship. Contractor personnel who furnish services under this contract are subject to Government technical oversight of the services provided. The Government retains the right to reject nonconforming services.

Contractors shall not perform inherently governmental functions including decision-making, supervision of Government employees, supervision of other Contractors on other contracts, and activities that create the appearance of performing such functions.

## Period of Performance

The anticipated period of performance of this task order (TO) is September 20, 2021 – September 19, 2022.

## Place of Performance

This is a nationwide contract to include Continental United States (CONUS) and Puerto Rico for all VA medical centers and outpatient clinics where Biomedical Engineering currently provides support.  Attachment A: List of Facilities identifies all main facilities where work may be performed as part of this contract.  The Contractor can determine how to service the outpatient clinics associated with main facilities.  The Government retains the right to add or subtract facilities as needed, and this will be done via modification to the contract.

## Hours of Performance

Normal working hours will range between 7:00 AM to 5:00 PM, Monday through Friday, excluding federal holidays, although there may be some working-hours flexibility.  This will be further defined in each task order. Occasional overtime may be needed and will be billed at the agreed rate. Contractor hours will be monitored by the L-COR and verified prior to billing and payment.

## Travel

Some tasks may require travel; authorized travel will be identified at the task order level. This travel must be pre-approved by the Local Contracting Officer’s Representative (L-COR) and will be reimbursed in accordance with Federal Travel Regulations (FTR) and Federal Acquisition Regulations (FAR) 31.205-46.   Each Contractor invoice shall include copies of all receipts that support the travel costs claimed in the invoice.  Local travel within a 50-mile radius from the duty location noted in the task order, is considered the cost of doing business and will not be reimbursed.  Travel performed for personal convenience and daily travel to and from work at the Contractor’s facility will not be reimbursed. General and Administrative (G&A) fees will not be reimbursed. If travel is not specifically authorized in writing by the L-COR in the applicable task order, it shall not be authorized as a separate billable expense.

## Scheduling and Availability Reports

The Contractor shall designate a Key Personnel for the entire contract as well as for each task order to address contract/task order needs. The Contractor shall provide an availability calendar sixty (60) days out indicating the number of qualified personnel available each week through the course of the year to aid in scheduling activities.

Deliverable: Availability calendar (due sixty (60) days after IDIQ Contract award)

### 7.1 Kick-Off Meeting:

The Contractor shall participate in Technical Kick-Off Meeting within ten calendar days after contract award.  The technical kick-off meeting shall occur in conjunction with the post award conference.  The Contractor shall present, for review and approval by the Government, the details of the intended approach, work plan, and projected schedule for this effort.  The Contractor shall provide an agenda at least five calendar days prior to the meeting, and document meeting minutes within three calendar days after the meeting.  The Contractor shall invite the Contracting Officer (CO), L-COR, Contract Specialist (CS), and the HTM/Biomedical Engineering’s operations staff.

Deliverable: Kickoff Meeting Agenda (five calendar days prior to the meeting) and Kick off Meeting Minutes (three calendar days after the kickoff meeting).

### 7.2 Monthly Status Reports:

The Contractor shall conduct monthly status report meetings with the facility-L-COR detailing, at a minimum, the previous month’s achievements, work planned for the current month, identified risks and recommendations to mitigate risk, staffing plans and forecasts, and a financial analysis of contract expenditures for the month (and year-to-date), including variance analysis compared to baseline and forecasted plans. Contractor monthly reports shall be submitted by first business day of each month.

Deliverable: Monthly Status Report (shall be submitted by first business day of each month to the L-COR by email)

END OF DELEVERABLES

## List of Equipment

All medical equipment listed in Attachment B: List of Types of Medical Equipment are included within the scope of this contract.  Medical equipment covered under the manufacturer’s warranty or covered under a service contract with another vendor is excluded from this contract.  The Government retains the right to add or subtract equipment as needed, and this will be done via a modification to the contract.

## Services Overview

At a high level, the tasks within the requirement include planned maintenance; corrective maintenance; and professional services for the technology assessment, implementation, and sustainment of medical devices and clinical systems, as well as healthcare consulting/internal operation support regarding medical equipment management.  The core functions are described in more detail as listed in Table 2: Core Biomedical Engineering Functions.

| Technology Assessment |  |
| --- | --- |
| TECHNOLOGY ASSESSMENT (PRE-PURCHASE) | This work is ordered and performed when local medical center leadership have a stated intent to acquire a medical device/system and extends through issuance of purchase order/contract for a recommended device/system. Technology Assessments include all work performed related to technology assessment and requirements definition prior to new equipment purchase order or contract award. - Support of Market Research: what features and functions are currently available; who are the manufacturer/suppliers; what are important design considerations; what the integration considerations are; cost considerations, support of IGCE formulation, etc. - Support of Requirements Definition: assistance determining “must have” and “desired” features/functions and level of technology; systems engineering design; support of writing technical specifications and RFP development - Support of usability testing and human factors design; assessing ease of use Contractor support for market research, requirements definition, and usability testing shall be limited to technical analysis and recommendations. Final decisions, approvals, and determinations shall be made by Government personnel. |
| SPACE AND UTILITIES CERTIFICATION | Space & utilities assessment and certification of availability or identification of needed space and utilities. Includes scoping and/or preliminary design of site prep construction associated with pre-purchase activities. For example, drafting the scope or preliminary design of a space to prepare it for medical equipment to be installed. |
| SECURITY RISK ASSESSMENT | All work performed prior to new equipment purchase order or contract award related to Pre- Procurement Risk Assessment for information security (e.g., Directive 6550, MDS2, etc.), data protection, and privacy. Includes assessment work regarding networked medical devices such as bandwidth and connectivity requirements for local area, wide area, and wireless networks. |
| Technology Implementation |  |
| SITE PREP SUPPORT | Work pertaining to site preparation support and space planning activities for the implementation of new equipment. This work generally follows issuance of purchase order/contract award for new medical equipment/systems and extends through the installation phase into initial clinical use. Includes site preparation planning, design, coordination, and execution, including collaboration with architects, engineers, equipment vendors, construction Contractors, and clinical departments. |
| PROJECT MANAGEMENT | Overall project management. Project management, coordination, and system configuration design with clinicians and vendors to implement new equipment/clinical systems. Includes preparation, participation, and leadership of meetings, conference calls, and communications. Includes creation, development, and upkeep of project documentation. This work is generally performed following issuance of purchase order or contract award for new equipment/systems and extends through first clinical use and transitions to operational sustainment. |
| PRE-IMPLEMENTATION SECURITY | Work associated with Information Security and Network Attached Medical Equipment Pre-Implementation Checklist (e.g., VA Directive 6550) and implementation of new equipment on medical Virtual Local Area Network (VLAN) with appropriate network isolation and protection. Includes design and configuration of Access Control Lists (ACL). This work is performed following issuance of purchase order or contract award for new equipment and generally extends through first clinical use and system optimization. |
| INSTALLATION | Installation work associated with implementation of new equipment/systems. Assembly and installation of new equipment items include hardware and software installation. Includes systems integration and connectivity testing. |
| INCOMING INSPECTION | Initial inspections and acceptance inspections to validate specifications and performance levels of new equipment, including owned, rental, loaner, and demo equipment. Includes safety inspection prior to first clinical use. This includes functional testing, user acceptance testing, and formal acceptance actions (per contract). Includes design and assignment of operations and maintenance strategy, developing Operations & Maintenance (O&M) plan and maintenance standard operating procedures. Establishing inventory record(s) and assigning maintenance schedule(s) for new equipment. |
| EDUCATION PROVIDED-IMPLEMENTATION/ USER | Provide or facilitate initial user training associated with new equipment implementation. Includes arranging vendor applications training and providing system “Go-Live” assistance and support. Includes training on operator-level maintenance and administration of equipment/system. |
| Technology Sustainment - Safety |  |
| HAZARD REMEDIATION | Actual corrective action and/or coordination of corrective actions for affected or recalled medical devices and associated items. |
| HAZARD/RECALL INVESTIGATION | Hazard/Recall review, discovery, and identification of affected devices and related reporting. Incident Investigation (separate from Information Security incidents): investigation of internal incidents to VA Medical Centers (VAMC(s)), including forensic activities, formal root cause analyses, Healthcare Failure Mode and Effects Analysis (HFMEA), and Safe Medical Device Act (SMDA)/MedWatch reporting, etc. |
| MEDICAL DEVICE SECURITY INCIDENT | Medical device security incident investigation, corrective action, and reporting/documentation. Includes responses to incidents such as malware intrusions, unauthorized/unsuccessful group policy pushes, and unauthorized/unexplained ACL change. |
| DE-INSTALLATION/ DECOMMISSIONING | Physical de-installation of a medical device/system. Decommissioning activities include sanitization of stored sensitive information as appropriate (e.g., physical destruction), reviewing maintenance history for completeness, assisting logistics with condition coding, removing from maintenance schedule, and other associated turn-in activities (abiding with Decommissioning standard operating procedures). |
| Technology Sustainment - Planned Maintenance |  |
| PLANNED MAINTENANCE | Routine support and planned maintenance of medical devices and medical systems (includes medical servers and other computer-based systems), including hardware, software, or communications networks. Typically performed according to defined procedures and on a scheduled basis, to minimize a device or system’s risk of failure and/or to ensure continued operation, safety, and data security. Includes tasks such as adjusting, calibrating, lubricating, cleaning, and replacing parts. Includes electrical safety inspection and testing. Includes testing the resiliency of medical systems infrastructure and fault tolerance testing. Includes visual inspection. Includes overhauls that are extensive replacing or rebuilding of system components (separate from upgrades). |
| ELECTRICAL SAFETY INSPECTION | Performance of electrical safety evaluations and applicable electrical measurements per VA, NFPA 99, and manufacturer specified requirements. |
| PERFORMANCE VERIFICATION | Verify device or system (includes medical servers and other computer-based hardware/software systems) “to be” operating state according to manufacturer and other specifications, separate from restorative repairs. Inspection to verify that the physical integrity, safety, and performance of a device/system meet necessary functional requirements. |
| MEDICAL DEVICE/SYSTEM PROTECTION | Planned/proactive inspection and installation of medical device operating systems, and malware prevention/antivirus software, and updates of same. Planned review of ACL design and update as appropriate; consideration of data backup and disaster recovery processes; test updating of automatic failover functions; assessment of data storage capacity and security of sensitive data. Performance of both physical and logical security reviews. Preparation of IT Requests for Medical Device Isolation Architecture (MDIA) actions, ACL, Changes, new Virtual Local Area Network (VLAN) assignments, and associated network jack installation/activation. Also includes OS, malware protection/antivirus software patching, and planned ACL and MDIA work activities. |
| Technology Sustainment - Corrective Maintenance |  |
| GENERAL REPAIR | Work performed on equipment to restore to clinical operation. Generally, hardware related. Restore physical integrity, safety, and/or performance of a device/system after a failure. This work typically includes troubleshooting and determination of the actions to complete the necessary repair(s); researching and determining necessary replacement parts; routine purchasing procedures; receipt of parts; completion of repair; and any intra-facility travel associated with pickup or return of equipment. Includes problem identification and resolution of medical systems servers that is primarily hardware related (may include incidental running of software routines, reloading software, or connectivity assessment that is incidental to the hardware resolution). |
| SOFTWARE REPAIR | Work performed to repair/restore medical device software or systems to clinical operation after a failure via software related actions including reloads, re-imaging, and other corrective actions. Includes problem identification and resolution of medical systems servers that is primarily software related (may include incidental troubleshooting and correction of hardware routines, reloading hardware, or connectivity assessment that is incidental to the software resolution). |
| MEDICAL SYSTEMS CONNECTIVITY | Determination of cause and correction of a reported operational problem(s) with a network-connected medical device or system determined to be related to network, interoperability protocol, or other connectivity-based infrastructure. Includes troubleshooting and correcting reported problems associated with Virtual Local Area Network (VLANs), ACLs, network port security, telecommunications, patch cable problems, Health Level 7 (HL7) outages, etc. Includes activities related to system connectivity or system wide related interoperability and interconnectivity issues. |
| Technology Sustainment |  |
| UPGRADES | Enhancement and upgrades for added functionality, greater reliability, and improved patient safety, including hardware and software improvement. Includes any significant reconfiguration of existing medical device/systems. |
| RELOCATION | Work associated with relocation of a medical device or system. |
| FABRICATION/ MODIFICATION | Work performed to design/create new equipment, assemblies, or apparatuses. Modification that alters a device/system to improve performance, reliability, or safety, or to add new functionality. |
| EDUCATION PROVIDED- CLINICAL USER | Providing training on the operation of medical equipment and technology systems. Includes training on operator-level maintenance and system administration. Includes time spent facilitating training with clinical educators, vendors, and other 3rd parties. Includes general training provided regarding medical equipment management, HTM services, equipment safety, new employee orientation, nurse staff meetings, etc., and mentoring of non-Biomedical Engineering personnel. |
| MEDICAL APPLICATION SUPPORT | Time spent on medical application changes. Work includes but not limited to adding/removing users, configuration changes, system application administration, regular medical device application software patching, clock resets, rebooting PCs, and interoperability verification. |
| Healthcare Consulting |  |
| EQUIPMENT PLANNING | Support of developing and monitoring VISN/VAMC strategic equipment plan(s). Includes supporting high-level needs assessment, technology research, business proposals, cost-benefit analyses, and lifecycle modeling. |
| LARGE PROJECT-CUSTOMER | Customer facing project/assignment/activity that takes more than 10 hours over the duration of the assignment. |
| GENERAL CONSULTATION | Technical consultation with clinical services, administrative services, research, executive leadership (VAMC, Veterans Integrated Service Network (VISN), and VHA/VA Program Offices) regarding healthcare technology matters. Includes serving on healthcare committees (VAMC and VISN levels) such as Critical Care, Radiation Safety, Laser Safety, Environment of Care, etc. as well as time-limited and ad-hoc teams such as patient care throughput/work, process flow, or other performance improvement team efforts. Includes time spent conducting “rounds” or other scheduled or unscheduled consultative visits with equipment users and customers. |
| MISCELLANEOUS | Other customer-facing services not specifically described in this table (e.g. responding to medical equipment-related questions from clinical staff, doing research on a medical equipment topic to assist another department, customer communications, etc.). |
| Internal Operations |  |
| PROJECT- INTERNAL | Project, activities, and assignments that are internal to HTM operations. |

Table 2: Core Biomedical Engineering Functions

### Planned Maintenance (PM)

Planned maintenance refers to a procedure ensuring proper medical equipment operation and calibration.  This procedure includes inspection of the medical equipment for mechanical defects, cleaning of interior fans and filters, performance verification of operator controls and adjustments, calibration verification per original equipment manufacturers’ (OEM) specifications, inspection for ground integrity and current leakage, and documentation of all results.  The Contractor shall have the appropriate level of training or OEM certification to perform PM services, and they shall not perform PM services on medical equipment that they are not fully trained or certified on.  The Contractor shall submit training certificates to the COR when required for the equipment being maintained.  The OEM’s specifications shall be used as the source document for all planned maintenance (unless otherwise instructed by the L-COR).

The Contractor shall ensure planned maintenance services support the VA facilities in preparation for The Joint Commission and other regulatory inspection surveys.  The L-COR will notify the Contractor by the end of the month prior that Contractor support is required for PM services in the following month.  The Contractor shall ensure all scheduled PM services are completed by the due date set in the VA Computerized Maintenance Management System (CMMS).  If at the due date, the Contractor still cannot complete the PM, all efforts to complete the PM shall be documented on a field service report and the information provided to the L-COR.  Contractor planned maintenance services shall include, but are not limited to the following:

- Performance verification

- Electrical safety inspection and testing

- Routinely scheduled service per OEM specifications

- Ensuring all equipment is maintained and repairing it to OEM standards if needed

- Documentation of all planned maintenance services upon completion or no later than the fifth (5th) business day of the subsequent month

### Corrective Maintenance

Medical equipment repairs shall be facilitated by the Contractor to ensure equipment availability, to provide an optimum level of equipment performance, and to extend the useful life of the equipment. The Contractor shall have the appropriate level of training or OEM certification to perform corrective maintenance services, and they shall not perform corrective maintenance services on medical equipment that they are not fully trained or certified on.  The Contractor shall submit training certificates to the COR when required for the equipment being repaired.  When repairs are required to correct equipment malfunctions, the Contractor shall use industry accepted repair practices and maintain high standards of equipment performance and safety to complete corrective maintenance services in a timely manner.  All efforts shall be made to meet an average corrective maintenance turn-around time of 7 calendar days from the initiation of a work order from the VA to the completion of the work order by the Contractor.  The Contractor corrective maintenance services shall include, but are not limited to the following:

- Performance analysis

- Troubleshooting and problem identification

- Decision-making regarding problem correction in the most pertinent, timely, and cost-effective manner

- Completing the repair

- Performance and electrical safety testing

- Documentation of all corrective maintenance services upon completion or within one (1) day of job completion.

### Professional Services

Professional services include any work related to supporting medical equipment or Biomedical Engineering department operations that is not planned or corrective maintenance.  The Contractor shall provide professional services in the areas of technology assessment, implementation, and sustainment of medical devices and clinical systems; healthcare consulting; and internal operation support.  Contractor services for Office of Electronic Health Record Modernization (OEHRM) deployments focus primarily on technology assessment, implementation, configuration management and healthcare consulting services.  Contractor services for operational support of a VA medical center or VISN could include any of the services listed.  Services breakdown will vary per location and Contractor skill level.  Contractor-provided professional services shall include, but are not limited to the following:

- Project Management - Overall project management. Project management, coordination, and system configuration design with clinicians and vendors to implement new equipment/clinical systems. Includes preparation, participation, and leadership of meetings, conference calls, and communications. Includes creation, development, and upkeep of project documentation. This work is generally performed following issuance of purchase order or contract award for new equipment/systems and extends through first clinical use and transitions to operational sustainment.

- Equipment and Systems Planning – Support assessing medical equipment and clinical systems portfolios, identify equipment approaching the end of its useful life, identify legacy systems, identify emerging or innovative technology/systems, support developing medical equipment and clinical systems strategic plans.

- Specification Development – Support defining and validating the requirements, compliance, features, and functions for medical equipment and clinical systems through extensive stakeholder interaction, technical subject matter expertise, and market research.

- Technical Analysis - Perform technical market analysis of medical equipment and clinical systems, considering clinical capabilities, technical performance, supportability, device integration, design and architecture, desired features, human factors design, ease of use, cost, and other factors.

- Risk Assessment - Perform risk assessments of medical equipment and clinical systems in collaboration with clinical stakeholders, patient safety personnel, and with IT for human factors design, reliability, information security, data protection, and connectivity requirements. Submit information required for Enterprise Risk Assessments (ERAs).

- Site Readiness - Collaborate with architects, engineers, equipment vendors, construction Contractors, clinical services, IT, and others to ensure proper medical equipment and clinical systems siting, safety compliance, physical and remote connectivity, security documentation, and code compliance in preparation for major medical equipment and clinical system installations.

- Installation and Inspection - Ensure equipment and systems are inspected and approved for implementation by validating specifications, performance levels, and electrical code compliance; performing functional testing, user acceptance testing, formal acceptance actions (per contract), and work associated with documentation; designing the equipment lifecycle strategies; developing preventative and corrective maintenance plan and standard operating procedures; facilitating proper user training; evaluating needs for system access and ensuring that the need to perform a task, permission levels, and competency requirements are met; ensuring proper nomenclature categorization; and assigning maintenance schedule(s) for new equipment and systems.

- Systems Design and Integration - Address and evaluate multiple parameters to ensure that separate medical equipment and clinical systems work as complete, integrated systems (when necessary).  Facilitate with Cerner-medical device integration parameter testing efforts, medical device configurations and Cerner Connectivity Engine (CCE) configurations required for Biomedical Device Integration (BMDI) solutions.

- Configuration - Arrange and document clinical system specifications, components, connectivity, and operating system and software versions; track and control changes; and validate and maintain consistency and completeness of the clinical system.

- Technical Training and Support - Provide or facilitate education and applications training for Biomedical Engineering staff and facilitate operator training around operating and maintaining medical equipment and clinical systems. Provide helpdesk and troubleshooting support for medical equipment and/or clinical systems.

- Safety Alerts, Recalls, and Incident Management - Manage, lead, and document activities associated with medical equipment and clinical systems safety, recalls, and incident investigations to prevent harm to patients, including proactive risk assessments, evidence preservation, data capture and analyses, root cause identification, corrective action development, solutions implementation, and disaster recovery. Set, monitor, and/or remediate clinical system alerts and failures, including performing remote monitoring and recovery.

- Cybersecurity Risk Management - Directly support and monitor medical equipment and clinical system cybersecurity compliance, including: ensuring proper medical network security operations, conducting planned and proactive updates of devices and systems, providing quality assurance on medical system compatibility, ensuring equipment inspections and calibrations do not degrade medical equipment performance, performing security incident investigation and corrective action, and testing device communication and system reporting to validate the expected level of integration and interoperability exists.

- Hardware and Software Upgrades - Proactively upgrade hardware and software to optimize patient safety and clinical efficacy through technology refresh cycles.

- Regulatory Compliance - Ensure compliance with VA and regulatory standards pertaining to medical equipment and clinical systems including those from VHA's Office of Healthcare Technology Management (HTM), Joint Commission (JC), National Fire Protection Association (NFPA), Food and Drug Administration (FDA), and other applicable federal, state, and local regulatory bodies.

- Data Analytics Support – Perform ad-hoc analytical Structured Query Language (SQL) queries, develop PowerBI reports, and develop SQL Server Reporting Services (SSRS) reports to support data analytics and data reporting for Biomedical Engineering operations.  Support ASP.NET MVC web development, troubleshoot analytics applications, and provide bug fixes to reports and web applications to assist national Biomedical Engineering data analytics initiatives.  Support in this area requires proficiencies in Microsoft PowerBI, Microsoft SQL Server Reporting Services (SSRS), ASP.NET MVC framework, and Web App Development and could require proficiencies in, C#, T-SQL, HTML and JavaScript, ETL Development using Microsoft SQL Server Data Tools (SSDT), and Internet Information Services (IIS).

### Remote Support

Remote Biomedical Engineering support is a valuable, feasible and timely alternative to onsite work in order to fulfill healthcare technology management needs.  Planned and corrective maintenance inherently cannot be performed remotely.  However, the Contractor may be required to perform professional services remotely, except for hardware and software upgrades, implementations, and technical training and support, which require onsite attendance.

### Test Equipment

The Contractor shall provide their own general test equipment to perform services covered under this contract, to include an electrical safety analyzer, gas flow analyzer, test lung simulator, defibrillator analyzer, pulse oximeter/ noninvasive blood pressure simulator analyzer, electrosurgical analyzer, external pacemaker analyzer, vital signs / patient ECG Simulator, bladder scanner phantom, x-ray test device, and network cable tester.  Any specialized test equipment owned by the VA can be used by the Contractor with approval from the L-COR. The Contractor shall notify the L-COR if any specialty test equipment is needed that is not available in-house, and if deemed necessary by the L-COR, the VA will procure the test equipment. The test equipment owned by the Contractor shall be on a calibration cycle, and the calibration of test equipment shall be traceable and in conformance with the test equipment OEM standards.  A copy of current test equipment calibration certificates used by the Contractor shall be provided by the Contractor at the beginning of the period of performance and at any time upon the request of the L-COR, COR or Contracting Officer.  All test equipment used for services rendered as part of this contract will be annotated by the L-COR on the VA’s work order including the certification expiration date for the test equipment.  Work completed by the Contractor using test equipment with expired calibration certificates will not be accepted.

### Parts and Service

The Contractor shall not be asked to work on medical equipment covered under the manufacturer’s warranty or covered under a service contract with another vendor.  If parts or medical equipment vendor services are needed for any work on medical equipment or accessories rendered as part of this contract, then the Contractor shall notify the L-COR of the exact parts or services that they require to complete the contracted work.  The L-COR will process the procurement of any parts or vendor services related to the work covered in this contract and will provide those parts or services to the Contractor for work completion as per the contract. NOTE: If services are required from the medical equipment OEM, or if the Contractor runs into an issue with a device that requires work from the medical equipment OEM of the medical device, then the Contractor shall inform the L-COR who will confirm the need and  procure those additional services from the OEM needed to complete the repair/maintenance on that device.

## Work Order Process

The method for which work orders will be initiated and tracked will be determined by the L-COR and followed by each applicable facility, but some acceptable examples include email, conference calls, Contractor-provided portal, or VA Computerized Maintenance Management System (CMMS).  Relevant work orders covered under this contract will be processed by the VA and the Contractor using the following steps:

- Initiation – The L-COR will initiate the work order to the Contractor through the L-COR determined and facility-determined method.  During this initiation, the L-COR will provide any necessary documentation for the Contractor to complete the work (e.g. Planned Maintenance (PM) procedures, issue descriptions, points-of-contact, etc.).  For the initiation of PM work orders, the L-COR will notify the Contractor by the end of the month prior that Contractor support is required for PM services in the following month.

- Confirmation – The Contractor shall confirm receipt of the work order within 1 business day.

- Scheduling – The Contractor shall work with the L-COR or Biomedical Engineering point-of-contact to schedule the work as soon as possible, taking into consideration the level of responsiveness needed, the type of medical equipment involved, and the resources that are available.

- Completion of Work – The Contractor shall notify the L-COR or Biomedical Engineering point-of-contact prior to commencing service during normal operating hours.  While on-site for maintenance, the Contractor shall be equipped with the appropriate test equipment based on the maintenance and repair contracted.  At that time the L-COR can request to see a copy of the test equipment calibration certificates to verify they are not expired.  If during the accomplishment of the work order, the Contractor required parts or vendor services to be procured, the Contractor shall notify the L-COR or Biomedical Engineering point-of-contact who will provide the parts or vendor services required for the Contractor to complete the contracted work.  The Contractor shall have a VA Contractor PIV card, and will be allowed to accomplish the work unescorted when in VA facilities.

- Close-Out – Once the work has been completed, the Contractor shall provide a field service report to the L-COR or Biomedical Engineering point-of-contact detailing the planned maintenance steps completed and the PM status or the cause of the equipment malfunction and corrective action that was taken.  The Contractor shall also include, in the field service report, the name of the person performing the work, the hours spent on the work order, the equipment ID#, and the equipment location.  All test equipment used for services rendered on a work order shall be annotated on the field service report, including the certification expiration date for the test equipment.  The Contractor shall provide the date the work was completed on the field service report, so that the L-COR or Biomedical Engineering point-of-contact can review, examine, determine acceptance, and close the work order with the appropriate date.

## Physical Demands

Work assignments will involve bench and on-site repairs requiring sitting, standing, walking, lifting, climbing, bending, and working in cramped and awkward positions.  Contractor personnel shall be able to frequently handle objects weighing up to 40 pounds and occasionally objects weighing more than 40 pounds.

## Work Environment

Work will be completed in the Biomedical Engineering department and patient care areas.  Work will primarily be done inside well-lit, heated, and ventilated areas, although access to some equipment may be more cramped and poorly lit.  Contractor personnel may be exposed to tremendously high electrical voltages and currents, and may sustain cuts, bruises or burns caused by electric shock, hot components, and ionizing radiation x-rays. Contractor personnel working near radiological equipment shall comply with VA radiation safety procedures. They may also be exposed to patient bodily fluids when working in patient care areas. Contractor personnel shall follow all applicable infection control protocols, including use of PPE, hand hygiene, and exposure reporting procedures, in accordance with VA facility policies. Please note that the Contractor shall provide their staff with Personal Protective Equipment (PPE) as VA is not liable for any injuries caused by a Contractor staff’s non-protection. Contractor personnel shall be familiar with VA facility emergency procedures including reporting incidents.

## Policies, Safety and Infection Control

The Contractor shall follow all relevant hospital and department-level policies and standard operating procedures (SOP’s) during performance of services under this contract.  Upon request from the Contractor, the L-COR will provide copies of any relevant hospital and department-level polices and SOP’s that the Contractor shall follow.

During the performance of services in this contract, the Contractor shall follow local safety and infection control processes.  These safety and infection control processes can include, but are not limited, to the following: ladder safety, laser safety, personal protective equipment, lock-out/tag-out, hazardous material disposal, infection control risk assessment (ICRA), and permitting.

## Contractor Key Personnel

### Basic Qualifications

The Contractor personnel shall meet the minimum following qualifications.

- Experience within a Biomedical Engineering department in a medical center

OR

- Understanding of VISN, VA medical center, and Biomedical Engineering department operations

### Role-Specific Qualifications

In addition to the basic qualifications, the Contractor personnel identified below are considered key personnel and shall meet the following role-specific qualifications based on the core functions and services each will be providing under this contract.

- Biomedical Technician / Equipment Support Specialist – General (Level I, II, III)

  - Education Requirement:

    - Associates or higher degree in the biomedical technology or biomedical engineering field; OR

    - Associates or higher degree in electronics, a field of information technology, or engineering (other than biomedical engineering) AND 1 year of creditable experience; OR

    - Completion of the Department of Defense (DOD) biomedical equipment technician training program; OR

    - Completion of a DOD training program in repair and maintenance of electronic systems or information technology AND 1 year of creditable experience; OR

    - Completion of a biomedical engineering technician or Healthcare Technology Management (HTM) certificate program that included an internship or apprenticeship; OR

    - Completion of biomedical engineering technician or HTM certificate program that did not include an internship AND 1 year of creditable experience; OR

    - Hold or having ever held certification as a Certified Biomedical Equipment Technician (CBET), Certified Radiology Equipment Specialist (CRES), Certified Laboratory Equipment Specialist (CLES), or Certified Healthcare Technology Manager (CHTM). Continuous maintenance of certification is not required throughout contract.

  - Demonstrated competency to perform operational verification procedures, planned maintenance, and repair service on various medical equipment.

  - Understanding of regulatory agencies’ requirements, industry regulatory requirements and International Standard Organization standards.

  - Familiarity with the setup and application of test equipment.

  - Familiarity of networking principles and language (e.g., Access Control List (ACL), Transmission Control Protocol/ Internet Protocol (TCP/IP), and Virtual Local Area Networks (VLAN))

  - Ability to read, analyze, and interpret technical literature, schematics, and drawings.

  - Level I Experience - A minimum of one year of hospital or healthcare experience performing medical equipment maintenance.

  - Level II Experience - A minimum of three years of hospital or healthcare experience performing medical equipment maintenance.

  - Level III Experience - A minimum of five years of hospital or healthcare experience performing medical equipment maintenance.

- Biomedical Technician / Equipment Support Specialist – Information Systems (e.g., MUSE, Intensive Care Unit Clinical Information System (ICU-CIS), Picture Archiving and Communication System (PACS), etc.)

  - Education Requirement:

    - Associates or higher degree in the biomedical technology or biomedical engineering field; OR

    - Associates or higher degree in electronics, a field of information technology, or engineering (other than biomedical engineering) AND 1 year of creditable experience; OR

    - Completion of the Department of Defense (DOD) biomedical equipment technician training program; OR

    - Completion of a DOD training program in repair and maintenance of electronic systems or information technology AND 1 year of creditable experience; OR

    - Completion of a biomedical engineering technician or Healthcare Technology Management (HTM) certificate program that included an internship or apprenticeship; OR

    - Completion of biomedical engineering technician or HTM certificate program that did not include an internship AND 1 year of creditable experience; OR

    - Hold or having ever held certification as a Certified Biomedical Equipment Technician (CBET), Certified Radiology Equipment Specialist (CRES), Certified Laboratory Equipment Specialist (CLES), or Certified Healthcare Technology Manager (CHTM). Continuous maintenance of certification is not required throughout contract.

  - A minimum of three years of hospital or healthcare experience performing medical equipment maintenance, with a minimum of one-year specialized experience in medical information systems and networking.

  - Current and advanced knowledge of computer virus protection software, available software patches and upgrades, information security tools, electronics, computers, and network theory.

  - Demonstrated competency in utilizing and maintaining computer and server-based medical equipment, medical record databases, and proprietary and generic software; experience with TCP/IP, HL7, networking, network security and Digital Imaging and Communications in Medicine (DICOM) standards; experience maintaining, interfacing and troubleshooting networked medical equipment to include VLAN configuration, IP addressing, sub-netting, and network security; ability to interpret, identify and apply network engineering principles and practices; and ability to distinguish networking problems from non-networking problems.

  - Understanding of regulatory agencies’ requirements, industry regulatory requirements and International Standard Organization standards.

  - Familiarity with the setup and application of network test equipment and tools.

  - Ability to read, analyze, and interpret technical literature, schematics, and drawings.

- Biomedical Technician / Equipment Support Specialist – Medical Imaging Systems (e.g. CT, MRI, X-Ray, Linear Accelerator)

  - Education Requirement:

    - Associates or higher degree in the biomedical technology or biomedical engineering field; OR

    - Associates or higher degree in electronics, a field of information technology, or engineering (other than biomedical engineering) AND 1 year of creditable experience; OR

    - Completion of the Department of Defense (DOD) biomedical equipment technician training program; OR

    - Completion of a DOD training program in repair and maintenance of electronic systems or information technology AND 1 year of creditable experience; OR

    - Completion of a biomedical engineering technician or Healthcare Technology Management (HTM) certificate program that included an internship or apprenticeship, OR

    - Completion of biomedical engineering technician or HTM certificate program that did not include an internship AND 1 year of creditable experience; OR

    - Hold or having ever held certification as a Certified Biomedical Equipment Technician (CBET), Certified Radiology Equipment Specialist (CRES), Certified Laboratory Equipment Specialist (CLES), or Certified Healthcare Technology Manager (CHTM). Continuous maintenance of certification is not required throughout contract.

  - A minimum of three years of hospital or healthcare experience performing medical equipment maintenance, with a minimum of one-year specialized experience with medical imaging equipment.

  - Comprehensive knowledge of imaging techniques, technologies, radiological practices, and procedures.

  - Demonstrated competency in maintaining medical imaging systems; ability to resolve highly complex breakdowns that have an immediate risk to patient safety; ability to maintain and troubleshoot networked medical imaging equipment and interfaces to include VLAN configuration, IP addressing, sub-netting, and network security, and build, configure, repair and install workstations and servers in imaging and non-imaging environments; ability to utilize and maintain computer and server based medical equipment, medical record databases, and proprietary and generic software; experience with TCP/IP, HL7, networking, and DICOM standards.

  - Understanding of regulatory agencies’ requirements, industry regulatory requirements and International Standard Organization standards.

  - Skill in using advanced radiological test equipment such as radiation dosimeters, calibration phantoms, high voltage testing devices, KV and MA meters and monitor calibration devices.

  - Ability to read, analyze, and interpret technical literature, schematics, and drawings.

- Biomedical Engineer / Clinical Engineer

  - A Bachelor’s degree or higher in Engineering or Certification as a Clinical Engineer (CCE).

  - A minimum of three years of hospital or healthcare experience maintaining medical equipment or managing a hospital medical equipment program.

  - Demonstrated knowledge of concepts related to computer based medical systems, networking protocols, and information security as it applies to medical technology within VHA.

  - Demonstrated competency in the ability to conduct a medical equipment management program that is compliant with applicable healthcare standards and regulatory agencies; ability to develop a curriculum for a continuing education program that address the safe and effective use of medical technology; ability to manage a recall and safety alert program for medical devices, including medical device hazard investigations, to assure compliance with patient safety goals, Safe Medical Device Act (SMDA), and JC requirements; ability to conduct capital asset and infrastructure planning for medical equipment spanning initial concept, installation, and effective implementation of complex medical equipment; ability to function as the subject matter expert in the field of biomedical engineering, directly supporting specialized clinical technology, including service, system administration, training, quality assurance, and life-cycle management; ability to effectively advise clinical and administrative staff on medical technology, including existing and emerging technology, which addresses viability, long-term suitability, compatibility, and/or safety; to apply project management principles to deployment of medical equipment and health information technologies.

  - Demonstrated skill in in communicating and working collaboratively with key stakeholders, including technical and professional staff at various levels of the organization.

  - Understanding of regulatory agencies’ requirements, industry regulatory requirements and International Standard Organization standards.

- Project Manager

  - Bachelor’s degree or higher and certification in project management (e.g., Project Management Professional (PMP)).

  - A minimum of two years of hospital or healthcare experience planning, managing, and controlling projects.

  - Demonstrated competency in planning, managing, and controlling projects throughout their lifecycle; writing or reviewing project management documentation (including project charters, project management plan, Business Reference Model (BRM); problem solving; and risk mitigation and management.

  - Demonstrated mastery and skill in applying the theory, concepts, practices, and techniques consistent with the project management principles outlined in the Project Management Institute's (PMI) Project Management Body of Knowledge (PMBOK).

  - Demonstrated skill in in communicating and working collaboratively with key stakeholders, including technical and professional staff at various levels of the organization.

  - Understanding of regulatory agencies’ requirements, industry regulatory requirements and International Standard Organization standards.

### Evidence

The Contractor shall furnish qualified personnel as required by contract specifications to accomplish services required under this contract.  The Contractor shall have evidence of training, certifications, experience, and other qualifications that is listed on their resume for their personnel performing work under the contract.  NOTE: There may be additional specific education or experience requirements at the task order level. Evidence shall include a list of the types of medical equipment on which contracted personnel have experience performing planned and corrective maintenance.  If the Contractor proposes to add or replace personnel to perform contract services, the Contractor shall submit the required evidence of training certification, experience, and other qualifications to the COR or L-COR at the task order level.  At no time shall the Contractor utilize additional or replacement personnel to perform contract services who do not meet the qualifications under the terms and conditions of this contract or have not been approved by the COR or L-COR at the task order level to provide support. Evidence can be subject to validation by the L-COR.

### Staffing Plan

Contractor shall provide resumes of proposed contract personnel. Selection of contract personnel to VA sites must be approved by designated VHA personnel before arrival and onboarding of credentials.

Contractor shall provide minimal notice of 72 hours to designated L-CORs of personnel removal or swap out/exchange for government to reassign work associated with contracted individuals.

Earliest billing of contract services can occur after creation of an active VA email and some method to access VA email (Virtual Desktop/Remote Access, etc).  Exemptions can be applied through the L-COR, if needed, depending on the situation presented (e.g., Contractor personnel working with a temporary ID while PIV is being processed).

Contractor shall assist contract personnel with navigating VHA onboarding process.

### Acceptance

Contractor personnel shall not be considered Government employees for any purpose. The VA reserves the right to restrict any Contractor personnel from performing services under this contract who do not meet the required qualifications for the services they are required to perform, and who violate Federal regulations or are identified as a potential threat to the security, safety, health and/or operational mission of VA and its’ Veteran population.  The restriction of such Contractor personnel shall not relieve the Contractor from performing all the required services, in accordance with all terms, conditions, and schedules contained herein.

## Badging and Security

Contractor personnel working under this contract are required to be fingerprinted and have background verification completed by the VSC Personnel Security Office for issuance of a VA Contractor PIV badge, prior to commencement of work (NOTE: A temporary ID badge may be provided prior to final verification and issuance of the PIV badge).  The Contractor PIV badge request must be initiated through the COR for this IDIQ contract or the L-COR for a specific TO.

Contractor personnel working under this contract are required to obtain a VHA government email account prior to commencement of work.

Contractor personnel working under this contract shall obtain VA GFE (Government Furnished Equipment) within 120 days of the commencement of work under this contract.

Contractor personnel providing networked medical system management services under this contract shall apply and be granted elevated privileges through Electronic Permission Access System (ePAS) within 150 days of the commencement of work under this contract.

All Contractor personnel working under this contract, requiring on-site access, shall always wear visible identification while on the premises of a VA facility. The Contractor shall park in the appropriate designated parking areas. Information on parking is available from the applicable VA Police Section or L-COR. The VA facility will not invalidate or make reimbursement for any parking violations of the Contractor under any conditions. Smoking is prohibited inside any building at the VAMC. Possession of weapons is prohibited. Enclosed containers, including tool kits, are subject to search. Violations of VA regulations will result in citations answerable in the United States (Federal) District court, not a local district, state, or municipal court.

## Contractor Performance

Contractor performance will be monitored by the L-COR.  Completed assignments will be reviewed and discussed with the L-COR. Incomplete or pending work orders, planned maintenance assignments, and other assigned tasks will be reviewed during weekly status meetings to ensure work is being completed in accordance with contract requirements.

The Government will evaluate the Contractor’s performance under this contract using the method of surveillance specified in Table 3: Quality Assurance Surveillance Plan.  All surveillance observations will be recorded by the Government through L-COR. When an observation indicates defective performance, the L-COR will provide the Contractor with a copy of the record of the observation.

| Paragraph Number | Performance Objective | Standard | Acceptable Quality Level (AQL) | Method of Surveillance |
| --- | --- | --- | --- | --- |
| 14 | Contractor Personnel Qualifications | The Contractor shall furnish qualified personnel as required by contract specifications to accomplish services under this contract. | At no time shall Contractor utilize additional or replacement personnel to perform contract services who do not meet the qualifications under the terms and conditions of this contract or have not been approved by the L-COR. | 100% Inspection |

| 9.1. and 9.2 | Planned Maintenance and Corrective Maintenance | Original Equipment Manufacturer Specifications | The Contractor Shall perform all maintenance and repair of equipment in accordance with original equipment manufacturer specifications 100% | Periodic Inspection |
| --- | --- | --- | --- | --- |

| 9.1., 9.2., and 10 | Service Reports | The Contractor shall document all work performed in accordance with Biomedical Standards Operating Procedures | 98% | 100% Inspection |
| --- | --- | --- | --- | --- |

Table 3: Quality Assurance Surveillance Plan

## Information Security/Privacy for Inclusion into Contracts

### 17.1	General

Contractors, Contractor personnel, subcontractors, and subcontractor personnel shall be subject to the same Federal laws, regulations, standards, and VA Directives and Handbooks as VA and VA personnel regarding information and information system security.

### 17.2	Access to VA Information and VA Information Systems

A Contractor/subcontrator shall request logical (technical) or physical access to VA information and VA information systems for their employees, subcontractors, and affiliates only to the extent necessary to perform the services specified in the contract, agreement, or task order.

All Contractors, subcontractors, and third-party servicers and associates working with VA information are subject to the same investigative requirements as those of VA appointees or employees who have access to the same types of information.  The level and process of background security investigations for Contractors must be in accordance with VA Directive and Handbook 0710, Personnel Suitability and Security Program.  The Office for Operations, Security, and Preparedness is responsible for these policies and procedures.

Contract personnel who require access to national security programs must have a valid security clearance.  National Industrial Security Program (NISP) was established by Executive Order 12829 to ensure that cleared U.S. defense industry contract personnel safeguard the classified information in their possession while performing work on contracts, programs, bids, or research and development efforts.  The Department of Veterans Affairs does not have a Memorandum of Agreement with Defense Security Service (DSS).  Verification of a Security Clearance must be processed through the Special Security Officer located in the Planning and National Security Service within the Office of Operations, Security, and Preparedness.

Custom software development and outsourced operations must be located in the U.S. to the maximum extent practical.  If such services are proposed to be performed abroad and are not disallowed by other VA policy or mandates, the Contractor/subcontractor must state where all non-U.S. services are provided and detail a security plan, deemed to be acceptable by VA, specifically to address mitigation of the resulting problems of communication, control, data protection, and so forth.  Location within the U.S. may be an evaluation factor.

The Contractor or subcontractor must notify the Contracting Officer immediately when an employee working on a VA system or with access to VA information is reassigned or leaves the Contractor or subcontractor’s employ.  The Contracting Officer must also be notified immediately by the Contractor or subcontractor prior to an unfriendly termination.

### 17.3	VA Information Custodial Language

- Information made available to the Contractor or subcontractor by VA for the performance or administration of this contract or information developed by the Contractor/subcontractor in performance or administration of the contract shall be used only for those purposes and shall not be used in any other way without the prior written agreement of the VA.  This clause expressly limits the Contractor/subcontractor's rights to use data as described in Rights in Data - General, FAR 52.227-14(d) (1).

VA information should not be co-mingled, if possible, with any other data on the Contractors/subcontractor’s information systems or media storage systems in order to ensure VA requirements related to data protection and media sanitization can be met.  If co-mingling must be allowed to meet the requirements of the business need, the Contractor must ensure that VA’s information is returned to the VA or destroyed in accordance with VA’s sanitization requirements.  VA reserves the right to conduct on site inspections of Contractor and subcontractor IT resources to ensure data security controls, separation of data and job duties, and destruction/media sanitization procedures are in compliance with VA directive requirements.

Prior to termination or completion of this contract, Contractor/subcontractor must not destroy information received from VA, or gathered/created by the Contractor in the course of performing this contract without prior written approval by the VA.  Any data destruction done on behalf of VA by a Contractor/subcontractor must be done in accordance with National Archives and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records and Information Management and its Handbook 6300.1 Records Management Procedures, applicable VA Records Control Schedules, and VA Handbook 6500.1, Electronic Media Sanitization.  Self-certification by the Contractor that the data destruction requirements above have been met must be sent to the VA Contracting Officer within 30 days of termination of the contract.

The Contractor/subcontractor must receive, gather, store, back up, maintain, use, disclose and dispose of VA information only in compliance with the terms of the contract and applicable Federal and VA information confidentiality and security laws, regulations and policies.  If Federal or VA information confidentiality and security laws, regulations and policies become applicable to the VA information or information systems after execution of the contract, or if NIST issues or updates applicable FIPS or Special Publications (SP) after execution of this contract, the parties agree to negotiate in good faith to implement the information confidentiality and security laws, regulations and policies in this contract.

The Contractor/subcontractor shall not make copies of VA information except as authorized and necessary to perform the terms of the agreement or to preserve electronic information stored on Contractor/subcontractor electronic storage media for restoration in case any electronic equipment or data used by the Contractor/subcontractor needs to be restored to an operating state.  If copies are made for restoration purposes, after the restoration is complete, the copies must be appropriately destroyed.

If VA determines that the Contractor has violated any of the information confidentiality, privacy, and security provisions of the contract, it shall be sufficient grounds for VA to withhold payment to the Contractor or third party or terminate the contract for default or terminate for cause under Federal Acquisition Regulation (FAR) part 12.

If a Veterans Health Administration (VHA) contract is terminated for default or cause with a business associate, the related local Business Associate Agreement (BAA) shall also be terminated and actions taken in accordance with VHA Directive 1605.05, Business Associate Agreements. If there is an executed national BAA associated with the contract, VA will determine what actions are appropriate and notify the contactor.

The Contractor/subcontractor must store, transport, or transmit VA sensitive information in an encrypted form, using VA-approved encryption tools that are, at a minimum, FIPS 140-2 validated.

The Contractor/subcontractor’s firewall and Web services security controls, if applicable, shall meet or exceed VA’s minimum requirements.  VA Configuration Guidelines are available upon request.

Except for uses and disclosures of VA information authorized by this contract for performance of the contract, the Contractor/subcontractor may use and disclose VA information only in two other situations: (i) in response to a qualifying order of a court of competent jurisdiction, or (ii) with VA’s prior written approval.  The Contractor/subcontractor must refer all requests for, demands for production of, or inquiries about, VA information and information systems to the VA contracting officer for response.

Notwithstanding the provision above, the Contractor/subcontractor shall not release VA records protected by Title 38 U.S.C. 5705, confidentiality of medical quality assurance records and/or Title 38 U.S.C. 7332, confidentiality of certain health records pertaining to drug addiction, sickle cell anemia, alcoholism or alcohol abuse, or infection with human immunodeficiency virus.  If the Contractor/subcontractor is in receipt of a court order or other requests for the above mentioned information, that Contractor/subcontractor shall immediately refer such court orders or other requests to the VA contracting officer for response.

For service that involves the storage, generating, transmitting, or exchanging of VA sensitive information but does not require C&A or an MOU-ISA for system interconnection, the Contractor/subcontractor must complete a Contractor Security Control Assessment (CSCA) on a yearly basis and provide it to the COTR.

### 17.4	Information System Design and Development

- Information systems that are designed or developed for or on behalf of VA at non-VA facilities shall comply with all VA directives developed in accordance with FISMA, HIPAA, NIST, and related VA security and privacy control requirements for Federal information systems.  This includes standards for the protection of electronic PHI, outlined in 45 C.F.R. Part 164, Subpart C, information and system security categorization level designations in accordance with FIPS 199 and FIPS 200 with implementation of all baseline security controls commensurate with the FIPS 199 system security categorization (reference Appendix D of VA Handbook 6500, VA Information Security Program).  During the development cycle a Privacy Impact Assessment (PIA) must be completed, provided to the COTR, and approved by the VA Privacy Service in accordance with Directive 6507, VA Privacy Impact Assessment.

The Contractor/subcontractor shall certify to the COTR that applications are fully functional and operate correctly as intended on systems using the VA Federal Desktop Core Configuration (FDCC), and the common security configuration guidelines provided by NIST or the VA.  This includes Internet Explorer 7 configured to operate on Windows XP and Vista (in Protected Mode on Vista) and future versions, as required.

The standard installation, operation, maintenance, updating, and patching of software shall not alter the configuration settings from the VA approved and FDCC configuration.  Information technology staff must also use the Windows Installer Service for installation to the default “program files” directory and silently install and uninstall.

Applications designed for normal end users shall run in the standard user context without elevated system administration privileges.

The security controls must be designed, developed, approved by VA, and implemented in accordance with the provisions of VA security system development life cycle as outlined in NIST Special Publication 800-37, Guide for Applying the Risk Management Framework to Federal Information Systems, VA Handbook 6500, Information Security Program and VA Handbook 6500.5, Incorporating Security and Privacy in System Development Lifecycle.

The Contractor/subcontractor is required to design, develop, or operate a System of Records Notice (SOR) on individuals to accomplish an agency function subject to the Privacy Act of 1974, (as amended), Public Law 93-579, December 31, 1974 (5 U.S.C. 552a) and applicable agency regulations.  Violation of the Privacy Act may involve the imposition of criminal and civil penalties.

The Contractor/subcontractor agrees to:

- Comply with the Privacy Act of 1974 (the Act) and the agency rules and regulations issued under the Act in the design, development, or operation of any system of records on individuals to accomplish an agency function when the contract specifically identifies:

- (a)  The Systems of Records (SOR); and

- (b)  The design, development, or operation work that the Contractor/subcontractor is to perform;

- Include the Privacy Act notification contained in this contract in every solicitation and resulting subcontract and in every subcontract awarded without a solicitation, when the work statement in the proposed subcontract requires the redesign, development, or operation of a SOR on individuals that is subject to the Privacy Act; and

- Include this Privacy Act clause, including this subparagraph (3), in all subcontracts awarded under this contract which requires the design, development, or operation of such a SOR.

In the event of violations of the Act, a civil action may be brought against the agency involved when the violation concerns the design, development, or operation of a SOR on individuals to accomplish an agency function, and criminal penalties may be imposed upon the officers or employees of the agency when the violation concerns the operation of a SOR on individuals to accomplish an agency function.  For purposes of the Act, when the contract is for the operation of a SOR on individuals to accomplish an agency function, the Contractor/subcontractor is considered to be an employee of the agency.

- “Operation of a System of Records” means performance of any of the activities associated with maintaining the SOR, including the collection, use, maintenance, and dissemination of records.

- “Record” means any item, collection, or grouping of information about an individual that is maintained by an agency, including, but not limited to, education, financial transactions, medical 
history, and criminal or employment history and contains the person’s name, or identifying number, symbol, or any other identifying particular assigned to the individual, such as a fingerprint or voiceprint, or a photograph.

- “System of Records” means a group of any records under the control of any agency from which information is retrieved by the name of the individual or by some identifying number, symbol, or other identifying particular assigned to the individual.

The vendor shall ensure the security of all procured or developed systems and technologies, including their subcomponents (hereinafter referred to as “Systems”), throughout the life of this contract and any extension, warranty, or maintenance periods.  This includes, but is not limited to workarounds, patches, hotfixes, upgrades, and any physical components (hereafter referred to as Security Fixes) which may be necessary to fix all security vulnerabilities published or known to the vendor anywhere in the Systems, including Operating Systems and firmware.  The vendor shall ensure that Security Fixes shall not negatively impact the Systems.

The vendor shall notify VA within 24 hours of the discovery or disclosure of successful exploits of the vulnerability which can compromise the security of the Systems (including the confidentiality or integrity of its data and operations, or the availability of the system).  Such issues shall be remediated as quickly as is practical, but in no event longer than ___7_ days.

When the Security Fixes involve installing third party patches (such as Microsoft OS patches or Adobe Acrobat), the vendor will provide written notice to the VA that the patch has been validated as not affecting the Systems within 10 working days.  When the vendor is responsible for operations or maintenance of the Systems, they shall apply the Security Fixes within ___7_ days.

All other vulnerabilities shall be remediated as specified in this paragraph in a timely manner based on risk, but within 60 days of discovery or disclosure.  Exceptions to this paragraph (e.g. for the convenience of VA) shall only be granted with approval of the contracting officer and the VA Assistant Secretary for Office of Information and Technology.

### 17.5	Information System Hosting, Operation, Maintenance, Or Use

- For information systems that are hosted, operated, maintained, or used on behalf of VA at non-VA facilities, Contractors/subcontractors are fully responsible and accountable for ensuring compliance with all HIPAA, Privacy Act, FISMA, NIST, FIPS, and VA security and privacy directives and handbooks.  This includes conducting compliant risk assessments, routine vulnerability scanning, system patching and change management procedures, and the completion of an acceptable contingency plan for each system.  The Contractor’s security control procedures must be equivalent, to those procedures used to secure VA systems.  A Privacy Impact Assessment (PIA) must also be provided to the COTR and approved by VA Privacy Service prior to operational approval.  All external Internet connections to VA’s network involving VA information must be reviewed and approved by VA prior to implementation.

Adequate security controls for collecting, processing, transmitting, and storing of Personally Identifiable Information (PII), as determined by the VA Privacy Service, must be in place, tested, and approved by VA prior to hosting, operation, maintenance, or use of the information system, or systems by or on behalf of VA.  These security controls are to be assessed and stated within the PIA and if these controls are determined not to be in place, or inadequate, a Plan of Action and Milestones (POA&M) must be submitted and approved prior to the collection of PII.

Outsourcing (Contractor facility, Contractor equipment or Contractor staff) of systems or network operations, telecommunications services, or other managed services requires certification and accreditation (authorization) (C&A) of the Contractor’s systems in accordance with VA Handbook 6500.3, Certification and Accreditation and/or the VA OCS Certification Program Office.  Government-owned (government facility or government equipment) Contractor-operated systems, third party or business partner networks require memorandums of understanding and interconnection agreements (MOU-ISA) which detail what data types are shared, who has access, and the appropriate level of security controls for all systems connected to VA networks.

The Contractor/subcontractor’s system must adhere to all FISMA, FIPS, and NIST standards related to the annual FISMA security controls assessment and review and update the PIA.  Any deficiencies noted during this assessment must be provided to the VA contracting officer and the ISO for entry into VA’s POA&M management process.  The Contractor/subcontractor must use VA’s POA&M process to document planned remedial actions to address any deficiencies in information security policies, procedures, and practices, and the completion of those activities.  Security deficiencies must be corrected within the timeframes approved by the government.  Contractor/subcontractor procedures are subject to periodic, unannounced assessments by VA officials, including the VA Office of Inspector General.  The physical security aspects associated with Contractor/subcontractor activities must also be subject to such assessments.  If major changes to the system occur that may affect the privacy or security of the data or the system, the C&A of the system may need to be reviewed, retested and re-authorized per VA Handbook 6500.3.  This may require reviewing and updating all of the documentation (PIA, System Security Plan, Contingency Plan).  The Certification Program Office can provide guidance on whether a new C&A would be necessary.

The Contractor/subcontractor must conduct an annual self assessment on all systems and outsourced services as required.  Both hard copy and electronic copies of the assessment must be provided to the COTR.  The government reserves the right to conduct such an assessment using government personnel or another Contractor/subcontractor.  The Contractor/subcontractor must take appropriate and timely action (this can be specified in the contract) to correct or mitigate any weaknesses discovered during such testing, generally at no additional cost.

VA prohibits the installation and use of personally-owned or Contractor/subcontractor-owned equipment or software on VA’s network.  If non-VA owned equipment must be used to fulfill the requirements of a contract, it must be stated in the service agreement, SOW or contract.  All of the security controls required for government furnished equipment (GFE) must be utilized in approved other equipment (OE) and must be funded by the owner of the equipment.  All remote systems must be equipped with, and use, a VA-approved antivirus (AV) software and a personal (host-based or enclave based) firewall that is configured with a VA-approved configuration.  Software must be kept current, including all critical updates and patches.  Owners of approved OE are responsible for providing and maintaining the anti-viral software and the firewall on the non-VA owned OE.

All electronic storage media used on non-VA leased or non-VA owned IT equipment that is used to store, process, or access VA information must be handled in adherence with VA Handbook 6500.1, Electronic Media Sanitization upon:  (i) completion or termination of the contract or (ii) disposal or return of the IT equipment by the Contractor/subcontractor or any person acting on behalf of the Contractor/subcontractor, whichever is earlier. Media (hard drives, optical disks, CDs, back-up tapes, etc.) used by the Contractors/subcontractors that contain VA information must be returned to the VA for sanitization or destruction or the Contractor/subcontractor must self-certify that the media has been disposed of per 6500.1 requirements.  This must be completed within 30 days of termination of the contract.

Bio-Medical devices and other equipment or systems containing media (hard drives, optical disks, etc.) with VA sensitive information must not be returned to the vendor at the end of lease, for trade-in, or other purposes.  The options are:

- Vendor must accept the system without the drive;

- VA’s initial medical device purchase includes a spare drive which must be installed in place of the original drive at time of turn-in; or

- VA must reimburse the company for media at a reasonable open market replacement cost at time of purchase.

- Due to the highly specialized and sometimes proprietary hardware and software associated with medical equipment/systems, if it is not possible for the VA to retain the hard drive, then;

- The equipment vendor must have an existing BAA if the device being traded in has sensitive information stored on it and hard drive(s) from the system are being returned physically intact; and

- Any fixed hard drive on the device must be non-destructively sanitized to the greatest extent possible without negatively impacting system operation.  Selective clearing down to patient data folder level is recommended using VA approved and validated overwriting technologies/methods/tools.  Applicable media sanitization specifications need to be pre-approved and described in the purchase order or contract.

- A statement needs to be signed by the Director (System Owner) that states that the drive could not be removed and that (a) and (b) controls above are in place and completed.  The ISO needs to maintain the documentation.

### 17.6	Security Incident Investigation

- The term “security incident” means an event that has, or could have, resulted in unauthorized access to, loss or damage to VA assets, or sensitive information, or an action that breaches VA security procedures.  The Contractor/subcontractor shall immediately notify the COTR and simultaneously, the designated ISO and Privacy Officer for the contract of any known or suspected security/privacy incidents, or any unauthorized disclosure of sensitive information, including that contained in system(s) to which the Contractor/subcontractor has access.

To the extent known by the Contractor/subcontractor, the Contractor/subcontractor’s notice to VA shall identify the information involved, the circumstances surrounding the incident (including to whom, how, when, and where the VA information or assets were placed at risk or compromised), and any other information that the Contractor/subcontractor considers relevant.

With respect to unsecured protected health information, the business associate is deemed to have discovered a data breach when the business associate knew or should have known of a breach of such information.  Upon discovery, the business associate must notify the covered entity of the breach.  Notifications need to be made in accordance with the executed business associate agreement.

In instances of theft or break-in or other criminal activity, the Contractor/subcontractor must concurrently report the incident to the appropriate law enforcement entity (or entities) of jurisdiction, including the VA OIG and Security and Law Enforcement.  The Contractor, its employees, and its subcontractors and their employees shall cooperate with VA and any law enforcement authority responsible for the investigation and prosecution of any possible criminal law violation(s) associated with any incident.  The Contractor/subcontractor shall cooperate with VA in any civil litigation to recover VA information, obtain monetary or other compensation from a third party for damages arising from any incident, or obtain injunctive relief against any third party arising from, or related to, the incident.

### 17.7	Liquidated Damages for Data Breach

Consistent with the requirements of 38 U.S.C. §5725, a contract may require access to sensitive personal information.  If so, the Contractor is liable to VA for liquidated damages in the event of a data breach or privacy incident involving any SPI the Contractor/subcontractor processes or maintains under this contract. However, it is the policy of VA to forgo collection of liquidated damages in the event the Contractor provides payment of actual damages in an amount determined to be adequate by the agency.

The Contractor/subcontractor shall provide notice to VA of a “security incident” as set forth in the Security Incident Investigation section above.  Upon such notification, VA must secure from a non-Department entity or the VA Office of Inspector General an independent risk analysis of the data breach to determine the level of risk associated with the data breach for the potential misuse of any sensitive personal information involved in the data breach.  The term 'data breach' means the loss, theft, or other unauthorized access, or any access other than that incidental to the scope of employment, to data containing sensitive personal information, in electronic or printed form, that results in the potential compromise of the confidentiality or integrity of the data.   Contractor shall fully cooperate with the entity performing the risk analysis.  Failure to cooperate may be deemed a material breach and grounds for contract termination.

Each risk analysis shall address all relevant information concerning the data breach, including the following:

- Nature of the event (loss, theft, unauthorized access);

- Description of the event, including:

- (a)  date of occurrence;

- (b)  data elements involved, including any PII, such as full name, social security number, date of birth, home address, account number, disability code;
     (3)  Number of individuals affected or potentially affected;

- (4)  Names of individuals or groups affected or potentially affected;

- (5)  Ease of logical data access to the lost, stolen or improperly accessed data in light of the degree of protection for the data, e.g., unencrypted, plain text;

- (6)  Amount of time the data has been out of VA control;
 (7)  The likelihood that the sensitive personal information will or has been compromised (made accessible to and usable by unauthorized persons);

- (8)  Known misuses of data containing sensitive personal information, if any;

- (9)  Assessment of the potential harm to the affected individuals;

- (10)  Data breach analysis as outlined in 6500.2 Handbook, Management of Security and Privacy Incidents, as appropriate; and

- (11)  Whether credit protection services may assist record subjects in avoiding or mitigating the results of identity theft based on the sensitive personal information that may have been compromised.

Based on the determinations of the independent risk analysis, the Contractor shall be responsible for paying to the VA liquidated damages in the amount of $37.50 per affected individual to cover the cost of providing credit protection services to affected individuals consisting of the following:

- (1)  Notification;

- (2)  One year of credit monitoring services consisting of automatic daily monitoring of at least 3 relevant credit bureau reports;

- (3)  Data breach analysis;

- (4)  Fraud resolution services, including writing dispute letters, initiating fraud alerts and credit freezes, to assist affected individuals to bring matters to resolution;

- (5)  One year of identity theft insurance with $20,000.00 coverage at $0 deductible; and

(6)  Necessary legal expenses the subjects may incur to repair falsified or damaged credit records, histories, or financial affairs.

### 17.8	Security Controls Compliance Testing

On a periodic basis, VA, including the Office of Inspector General, reserves the right to evaluate any or all of the security controls and privacy practices implemented by the Contractor under the clauses contained within the contract.  With 10 working-day’s notice, at the request of the government, the Contractor shall fully cooperate and assist in a government-sponsored security controls assessment at each location wherein VA information is processed or stored, or information systems are developed, operated, maintained, or used on behalf of VA, including those initiated by the Office of Inspector General.  The government may conduct a security control assessment on shorter notice (to include unannounced assessments) as determined by VA in the event of a security incident or at any other time.

### 17.9	Training

- All Contractor employees and subcontractor employees requiring access to VA information and VA information systems shall complete the following before being granted access to VA information and its systems:

- Sign and acknowledge (either manually or electronically) understanding of and responsibilities for compliance with the  Contractor Rules of Behavior, Appendix E relating to access to VA information and information systems;

- Successfully complete the VA Cyber Security Awareness and Rules of Behavior training and annually complete required security training;

- Successfully complete the appropriate VA privacy training and annually complete required privacy training; and

- Successfully complete any additional cyber security or privacy training, as required for VA personnel with equivalent information system access [to be defined by the VA program official and provided to the contracting officer for inclusion in the solicitation document – e.g., any role-based information security training required in accordance with NIST Special Publication 800-16, Information Technology Security Training Requirements.]

The Contractor shall provide to the contracting officer and/or the COTR a copy of the training certificates and certification of signing the Contractor Rules of Behavior for each applicable employee within 1 week of the initiation of the contract and annually thereafter, as required.

Failure to complete the mandatory annual training and sign the Rules of Behavior annually, within the timeframe required, is grounds for suspension or termination of all physical or electronic access privileges and removal from work on the contract until such time as the training and documents are complete.

## Records Management Language for Contracts

- Contractor must comply with all applicable records management laws and regulations, as well as National Archives and Records Administration (NARA) records policies, including but not limited to the Federal Records Act (44 U.S.C. chs. 21, 29, 31, 33), NARA regulations at 36 CFR Chapter XII Subchapter B, and those policies associated with the safeguarding of records covered by the Privacy Act of 1974 (5 U.S.C. 552a). These policies include the preservation of all records, regardless of form or characteristics, mode of transmission, or state of completion.

- In accordance with 36 CFR 1222.32, all data created for Government use and delivered to, or falling under the legal control of, the Government are Federal records subject to the provisions of 44 U.S.C. chapters 21, 29, 31, and 33, the Freedom of Information Act (FOIA) (5 U.S.C. 552), as amended, and the Privacy Act of 1974 (5 U.S.C. 552a), as amended and must be managed and scheduled for disposition only as permitted by statute or regulation.

- In accordance with 36 CFR 1222.32, Contractor must maintain all records created for Government use or created in the course of performing the contract and/or delivered to, or under the legal control of the Government and must be managed in accordance with Federal law. Electronic records and associated metadata must be accompanied by sufficient technical documentation to permit understanding and use of the records and data.

- The Government agency and its Contractors are responsible for preventing the alienation or unauthorized destruction of records, including all forms of mutilation. Records may not be removed from the legal custody of the Government agency or destroyed except for in accordance with the provisions of the agency records schedules and with the written concurrence of the Head of the Contracting Activity. Willful and unlawful destruction, damage or alienation of Federal records is subject to the fines and penalties imposed by 18 U.S.C. 2701.

In the event of any unlawful or accidental removal, defacing, alteration, or destruction of records, Contractor must report to the Government agency. The agency must report promptly to NARA in accordance with 36 CFR 1230.

- The Contractor must immediately notify the appropriate Contracting Officer upon discovery of any inadvertent or unauthorized disclosures of information, data, documentary materials, records or equipment. Disclosure of non-public information is limited to authorized personnel with a need-to-know as described in the [contract vehicle]. The Contractor must ensure that the appropriate personnel, administrative, technical, and physical safeguards are established to ensure the security and confidentiality of this information, data, documentary material, records and/or equipment is properly protected. The Contractor must not remove material from Government facilities or systems, or facilities or systems operated or maintained on the Government’s behalf, without the express written permission of the Head of the Contracting Activity. When information, data, documentary material, records and/or equipment is no longer required, it must be returned to the Government agency’s control or the Contractor must hold it until otherwise directed. Items returned to the Government must be hand carried, mailed, emailed, or securely electronically transmitted to the Contracting Officer or address prescribed in the [contract vehicle]. Destruction of records is EXPRESSLY PROHIBITED unless in accordance with Paragraph (4).

- The Contractor is required to obtain the Contracting Officer's approval prior to engaging in any contractual relationship (sub-Contractor) in support of this contract requiring the disclosure of information, documentary material and/or records generated under, or relating to, contracts. The Contractor (and any sub-Contractor) is required to abide by Government agency guidance for protecting sensitive, proprietary information, classified, and controlled unclassified information.

- The Contractor must only use Government IT equipment for purposes specifically tied to or authorized by the contract and in accordance with Government agency policy.

- The Contractor must not create or maintain any records containing any non-public Government agency information that are not specifically tied to or authorized by the contract.

- The Contractor must not retain, use, sell, or disseminate copies of any deliverable that contains information covered by the Privacy Act of 1974 or that which is generally protected from public disclosure by an exemption to the Freedom of Information Act.

- The Government agency owns the rights to all data and records produced as part of this contract. All deliverables under the contract are the property of the U.S. Government for which the Government agency must have unlimited rights to use, dispose of, or disclose such data contained therein as it determines to be in the public interest. Any Contractor rights in the data or deliverables must be identified as required by FAR 52.227-11 through FAR 52.227-20.

- Training.  All Contractor employees assigned to this contract who create, work with, or otherwise handle records are required to take VHA-provided records management training, Talent Management System (TMS) Item #10176, Privacy and Information Security, Rules of Behavior. The Contractor is responsible for confirming training has been completed according to agency policies, including initial training and any annual or refresher training.

## Attachment A: List of Facilities

## Attachment B: List of Types of Medical Equipment

## Attachment E: Pricing Spreadsheet