Section A: Midmark Delivery Systems (Brand Name Only)

Justification: Must physically mount to, and mechanically/electrically integrate with, the Government's existing Midmark dental chairs without voiding warranties or requiring unauthorized modifications.

| Item | Requirement | Salient Characteristics / Minimum Specifications |
| --- | --- | --- |
| 1 | Midmark ASepsis 21 Continental Radius Delivery system | - Must be compatible with the existing Midmark chair radius mount. - Must allow for left or right-handed operational conversion. - Required configuration (Left to Right): Air/Water (A/W) syringe, one (1) open position, one (1) 6-hole fiber optic connection, one (1) 6-hole fiber optic connection. |
| 2 | Midmark ASepsis 21 Assistant's Delivery system | - Must be a rear-mount, adjustable arm designed for Midmark chairs. - Must include mounting hardware for chair attachment. - Must be equipped with: one (1) High Volume Evacuator (HVE) Valve, one (1) saliva ejector, and one (1) A/W syringe. |
| 3 | Midmark HP Illumination system | - Must integrate seamlessly with the ASepsis 21 delivery unit. - Must provide reliable fiber optic power/illumination to the two 6-hole fiber optic handpiece positions specified in Item 1. |
| 4 | ISO-C F.O. Handpiece Tubing | - Must be ISO-C standard compliant. - Must support 6-pin fiber optic handpiece configurations. - Must integrate with the Midmark delivery system routing. |

Section B: Sterilization & Maintenance Equipment (Brand Name or Equal)

| Item | Requirement | Salient Characteristics / Minimum Specifications |
| --- | --- | --- |
| 5 | SciCan Statim G4 5000 Sterilizer (or Equal) | - Must be a cassette-style autoclave sterilizer. - Must have a rapid sterilization cycle (wrapped cycle time of approximately 17.5 minutes or less). - Must feature automated data logging (e.g., WiFi/LAN connectivity to network, print, or store cycle data digitally). - Must utilize single-use water system or equivalent to prevent cross-contamination of instruments. |
| 6 | Automatic Dental Handpiece Lubrication System | - Must automate the purging, lubricating, and maintenance processes for dental handpieces. - Must process multiple handpieces simultaneously per cycle. - Critical: Must be natively compatible, or include necessary adapters, to securely connect and service both J. Morita and KaVo brand handpieces. |

Section C: Endodontic Specialty Equipment (Brand Name or Equal)

| Item | Requirement | Salient Characteristics / Minimum Specifications |
| --- | --- | --- |
| 7 | J. Morita ZX2 Apex Locator (or Equal) | - Must provide highly accurate electronic root canal length measurement. - Must function accurately in both wet and dry canal conditions (blood, saline, bleach). - Must feature a clear visual display (LCD/Color) for real-time file tracking. |
| 8 | J. Morita Tri-Auto ZX2T (or Equal) | - Must be a cordless, handheld endodontic motor. - Must feature an integrated/built-in apex locator. - Must include automatic safety functions such as auto-reverse or torque-reduction (e.g., Optimum Torque Reverse) to prevent file separation. |

Section D: High & Low Speed Handpieces (Brand Name or Equal)

| Item | Requirement | Salient Characteristics / Minimum Specifications |
| --- | --- | --- |
| 9 | KaVo Expert Torque mini-Lux E675L (or Equal) | - High-speed, air-driven handpiece. - Must feature a "mini" head design for improved pediatric or posterior access. - Must include fiber optic illumination. - Must utilize ceramic bearings for durability. - Must be compatible with MULTIflex-style quick disconnect couplers. |
| 10 | KaVo Multiflex Lux Coupler 465LED (or Equal) | - Must feature a 6-pin connection to the delivery unit tubing. - Must include an integrated LED light source. - Must provide 360-degree swivel rotation to reduce hand fatigue. - Must include an anti-retraction valve to prevent cross-contamination. |
| 11 | KaVo Mastermatic Lux M20L Slowspeed (or Equal) | - Slow-speed contra-angle handpiece attachment. - Must have a 1:1 gear transmission ratio. - Must feature fiber optic illumination. - Must utilize a latch chuck mechanism. - Must connect via standard E-type connection. |
| 12 | KaVo Low Speed Air Motor 181L (or Equal) | - Must be a low-speed air motor (approximate speed range of 5,000 to 20,000 RPM). - Must be compatible with standard 4-hole or 6-pin tubing connections. - Must provide internal water/air coolant lines and support fiber optic illumination to attachments. |