| SUBJECT* | CTX Lensometer |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 76504 |
| --- | --- |
| SOLICITATION NUMBER* | 36C25726Q0538 |
| RESPONSE DATE/TIME/ZONE | 05-13-2026 10AM CENTRAL TIME, CHICAGO, USA |
| ARCHIVE | 30 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | 6515 |
| NAICS CODE* | 339115 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Network Contracting Office 17 Network Contracting Office 17 San Antonio TX 78240 |
| POINT OF CONTACT* | Contract Specialist Michelle Cunningham michelle.cunningham@va.gov 254-408-1652 |

PLACE OF PERFORMANCE

| ADDRESS | Central Texas VA Healthcare System |
| --- | --- |
|  | Olin E. Teague Medical Center |
|  | 1901 Veterans Memorial Drive |
|  | Temple Texas |
| POSTAL CODE | 76504 |
| COUNTRY | USA |

ADDITIONAL INFORMATION

| AGENCY’S URL | https://www.va.gov |
| --- | --- |
| URL DESCRIPTION | VA |
| AGENCY CONTACT’S EMAIL ADDRESS | michelle.cunningham@va.gov |
| EMAIL DESCRIPTION | POC |

DESCRIPTION

This market research notice is being conducted by Network Contracting Office (NCO) 17 VHA Regional Procurement Office-West (RPOW) 5441 Babcock Road Medical Center. 302 San Antonio TX 78240 for the procurement of for Ophthalmology department Lensometer.

- Background and Requirement: Central Texas Veterans Healthcare System, Olin E. Teague Medical Center located at 1901 Veterans Memorial Drive, Temple, Texas 76504 .This will be used to accurately measure the dioptric power, axis, cylinder, and prism of prescription lenses, including eyeglasses and contact lenses. It serves to verify that lenses match the doctor's prescription, locate optical centers, and mark lenses for proper fitting. These technologies are essential for guiding treatments and ensuring accurate diagnoses in eye care

- Ophthalmology department Service requires twenty-three (23) Lensometers  this equipment is used to accurately measure the dioptric power, axis, cylinder, and prism of prescription lenses, including eyeglasses and contact lenses.

- It serves to verify that lenses match the doctor's prescription, locate optical centers, and mark lenses for proper fitting. These technologies are essential for guiding treatments and ensuring accurate diagnoses in eye care.

- This equipment must be able to perform the following functions for all the SPECIFIC INSTRUMENT REQUIRED: Central Texas Veterans Healthcare System, Olin E. Teague Medical Center, Ophthalmology department requires 23 Lensometers.

- If this is an item your business can provide, please respond to this Sources Sought Notice by 10:00 AM Central Time on Wednesday 13 May 2026.

- Technical Requirements:

Measurement range

- Sphere (Spectacle lenses) -25 to +25 D

- Sphere (Contact lenses) -25 to +25 D (BC=6.0 to 9.0)

- (0.01 / 0.06 / 0.12 / 0.25 D increments)

- Cylinder 0 to ±10 D (-, MIX, +)

- (0.01 / 0.06 / 0.12 / 0.25 D increments)

- Axis 0 to 180° (1° increments)

- ADD 0 to +10 D Add, 2nd Add 2 (trifocal) Add 2 (trifocal)

- (0.01 / 0.06 / 0.12 / 0.25 D increments)

- Prism 0 to 20s

- (0.01 / 0.06 / 0.12 / 0.25s increments)

- Prism mode, Base In / Out, Base Up / Down; Polar Coordinates

- PD measurement 20.0 to 49.5 mm (monocular), Single vision PD, Progressive lens far vision PD

- UV transmittance 0 to 100% (1 or 5% increments)

- Measuring time 0.06 second ±10% (minimum)

- Measurable lens diameter

- Spectacle lenses 20 to 120 mm

- Contact lenses Larger than the inner diameter of the nosepiece within an aperature of ø5 mm

- Measurable transmittance 10% and over (20% and over for ±15 to ±25 D)

- 19. Compensation function for high index lenses: The abbe number is changeable in the range of 20   to 60.

- Marking system Ink cartridge type (Red)

- Wavelength / Measuring points 535 nm (green) / 108 within 7.5mm nosepiece

- Display 5.7-inch color full graphic TFT-LCD, 640 x 480 dots with LED backlight

- Printer Thermal line printer with auto cutter (paper width: 58 mm)

- Interface RS-232C (in / out), LAN, USB, Eye Care IC Card System

- Power Supply AC 100 to 240 V, 50 / 60 Hz

- Power consumption 60 VA

- Dimensions / Mass 220 (W) x 252 (D) x 430 (H) mm / 5.0 kg

- 8.7 (W) x 9.9 (D) x 16.9 (H) “ / 11.0 lbs.

- Standard accessories Printer paper, Power cord, Dust cover, Nosepiece for

- contact lenses, Explanation Guide for Measuring Progressive Power Lenses

- Optional accessories Eye Care cards, Interface cable, USB cable, Foot switch, Ink cartridges (Blue)

- Barcode scanner, Magnetic card reader.

a. The intent of this notice is to establish sources to define the procurement strategy (e.g. set-aside, sole source, unrestricted) for a solicitation that VA intends to post soon.  Interested contractors are requested to respond in accordance with the following:

- Where is the item manufactured, the VA is interested in procuring items manufactured in the United States. Provide City and State.

- If you have an existing GSA or VA, Federal Supply Schedule contract, include the contract details in your response.

- Please note that VA is particularly interested in determining the availability of Service Disable Veteran Owned Small Business/Veteran Owned Small Business or other Small Business manufacturers or authorized distributors that can provide this item.

- Respond to this notice if you can provide these products listed in the table in the background and requirement section above. In response, please cite your business size status.

- DISTRIBUTORS:  You must provide proof that you are an authorized distributor. VA does not accept gray market items.

- Vendors are requested to submit estimated market research pricing with their responses.  The estimated pricing will be considered when determining the procurement strategy for the future solicitation. (e.g., if CO determines that capable small businesses cannot provide fair and reasonable pricing, then the solicitation will not be set-aside).

- Responses should include:

(1) Business Name

(2) Business Address

(3) Point of Contact

- Name

- Phone Number

- E-mail Address

(4) SAM UID Number

(5) NAICS code

(6) Business Size: SMALL or LARGE

(7) Type of Business: SERVICE-DISABLED VETERAN OWNED (SDVOSB), VETERAN OWNED(VOSB), 8A, HUBZone, WOMAN-OWNED (WOSB), etc.

(8) Letter stating OEM authorized dealer, authorized distributor, or authorized reseller for the proposed equipment and/or services if applicable

(9) GSA/FSS Contract Number if applicable

All businesses eligible to provide this item are encouraged to reply as this information may be used to determine potential set-asides for the above-noted requirement.

8. Send responses to Michelle Cunningham at michelle.cunningham@va.gov.

DISCLAIMER:  THIS NOTICE OF INTENT IS NOT A REQUEST FOR COMPETITIVE QUOTES. However, Contractors may submit a capability statement, proposal, or quotation which shall be considered by the contracting officer whether to continue with the sole source procurement or to provide for competition. Due to technical issues, please allow additional time if you’re using electronic means.

Responses received will be evaluated; however, a determination by the Government not to compete with the proposed procurement based upon responses received to this notice is solely within the discretion of the Government.