-Provide analyzers that are FDA approved, 100-240V, and use FDA approved software for FDA approved testing reagents.

-Provide medium to high-capacity analyzer that can result at least 300 slides or higher per hour.

-Analyzer must have automated programming start, testing and shutdown capability.

-Analyzer should perform continuous loading and batch loading.

-Analyzer should perform automated staining and automated drying programming protocols.

-Analyzer must have barcode reader and disk drive with read and write and USB port for loading and downloading data and on-board operator's/ troubleshooting manuals.

-Must include installation of tubing or any other supplies required for water source or waste drain for operation.

-Analyzer Operating system must be compatible with Windows 10 or higher.

-Provides analyzer that interfaces with MHS Genesis laboratory information system (LIS) Cerner and anatomical pathology laboratory information system Pathnet and CoPath Plus.

-Provide analyzer that can interface with a specimen tracking system.

-Analyzer must provide capability for data and information storage and data management including maintenance.

-Provide on-site technical support and all reagents and supplies required for installation, validation/correlation of analyzers to bring to a fully operational status is included.

-Provide equipment repairs scheduled and unscheduled as well as preventive maintenance.

-Provide 24-hour telephonic troubleshooting.

-Respond to service calls within four hours or the morning of the next day of the next duty if after 5:00pm Monday-Friday and correct deficiencies within 24 hours after arrival on-site.

-Provide temporary analyzer until the original analyzer is repaired or permanently if the original analyzer can’t be repaired within 10 days.

-Provide alternate testing and/or validation for extended analyzer downtime/ failure.

-Provide on-site technical support for installation of analyzers to bring to and maintain at fully operational status including required reagents and supplies.

-Provide initial training at vendor’s facility for two personnel and on-site training as needed.

-Ensure proper packaging of reagents, supplies, and equipment and that they are shipped at proper refrigerated or frozen temperatures according to manufacturer environmental specifications.

-Provide FDA approved reagents for use and expiration dates of no less than 6 months from delivery date.

-Provide ready to use reagents that do not have to be prepared.

-Ensure all shipments arrive Monday through Friday between the hours of 8:00am and 3:00pm EST; excluding holidays and weekends.

-Respond to phone calls regarding shipment discrepancies the same day of notification and discrepancy will be corrected within 24 hours.

-Notify the laboratory if there is a recall on any reagents or equipment within 24 hours.

-Provide customer notification of updates on any reagents in writing within 5 business days and ensure continuity of services and availability of reagents during reagent revision when required.

--Provide all software and part upgrades at no additional cost.

-Must be able to start and obtain DOD Risk Management Framework (RMF) IT cybersecurity certification requirements.

-Analyzer must include online venting for fumes and filters.

-Analyzer must be able to provide the following: Hematoxylin and Eosin staining, special stains, PAP staining and film cover slipping capabilities.