| SUBJECT* | WLR Leica DM2700 P for TL Specifically configured for Asbestos Analysis (VA-25-00098025) |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 53214-1476 |
| --- | --- |
| SOLICITATION NUMBER* | 36C25226Q0295 |
| RESPONSE DATE/TIME/ZONE | 04-30-2026 11:00 am CENTRAL TIME, CHICAGO, USA |
| ARCHIVE | 60 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| PRODUCT SERVICE CODE* | 6515 |
| NAICS CODE* | 334516 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs 36C252 115 South 84th Street, Suite 101 Milwaukee WI 53214-1476 |
| POINT OF CONTACT* |  |

ADDITIONAL INFORMATION

| AGENCY’S URL |  |
| --- | --- |
| URL DESCRIPTION |  |
| AGENCY CONTACT’S EMAIL ADDRESS |  |
| EMAIL DESCRIPTION |  |

DESCRIPTION

Q&A

To clarify Question Number 11 from the previous modification Q&A.

11. Will a factory calibration certificate with NIST-traceable micrometer scale satisfy the requirement, or is third-party calibration preferred?

Unclear what requirement is being referred to here. But Factory calibration certification with the NIST traceable micrometer scale is a requirement that is requested.

Response

We can provide a NIST traceable calibration, but we need something to calibrate it to, and since there is no camera or measuring reticle on the proposal, I need guidance on what we will be calibrating and reporting the micrometer to.

I think the suggesting to test the system similar to the NIOSH 7400 method, against a certified phase-contrast test slide.