Minimum Essential Characteristics (MECs)

FORMULATRIX Flow Axial Seal Tip (F.A.S.T.) Liquid Handler

- Shall be no more than 532mm (W) x 403mm (H) x 596mm (D) in dimensions.

- Shall not weight over 18kg.

- Shall have a small footprint suitable for desktop use but is also open for integration into larger automated systems.

- Allows for rapid and efficient transfer of liquids in 96-well or 384-well plates, and other formats.

- Capable of dispensing volumes from 0.1 µL to 13 µL with high precision and accuracy, minimizing reagent usage. Transfer precision of 5% coefficient of variation (CV) at 0.1 µL.

- Shall offer both low volume (0.1-13 µL) and high volume (0.5-50 µL) tips, compatible with various plate types on a single head unit.

- Shall have multiple transfer modes allowing to support tasks like simple transfers, plate copies, cherry-picking, serial dilutions, and more.

- Shall include tip position sensors, automatic tip arrangement, and collision detection for reliable and accurate transfers.

- Shall offer an open Application Programming Interface (API) for easy integration with other laboratory automation systems.

- Shall utilize positive displacement disposable tips to ensure the integrity of samples without any risk of cross-contamination.

- Shall provide contact dispensing and have wide viscosity range.

- Shall allow for interchangeable decks allow for set up of tips, plates, and reagents prior to working on the instrument.

- Shall allow for semi or fully automated labware definitions and calibration.

- Shall allow for automated audit trails.

- Shall include Trusted Platform Module (TPM) 2.0 compatible Debian Linux on board computer.

- Shall allow for web browser-based operational controls such as Google Chrome version 91 or higher or other compatible internet browser interface.

- Shall be compatible with Integrated Fluidic Circuits (IFCs) 192.24, 96.96, 48.84.

- Shall be a validated solution for Olink IFCs and Proximity Extension Assay (PEA) Target and Explore workflows including Target 48 and 96, Explore 3072, and Explore-High Throughput (HT) kits.

- Shall allow for both Olink pre and post workflows.

- Shall include Olink Protocol bundles, Installation Qualification (IQ), Operational Qualification (OQ), and new install training.

- Shall be shipped and delivered to Robert Grant Ave, BLDG 503-Ground West 47,
Silver Spring, MD 20910-7500