# PRODUCT DESCRIPTION (PD)

# DEPARTMENT OF VETERANS AFFAIRS

EL Paso TX VHA Station 756

Biomedical Support Services

IT Support for GE MUSE Application Servers with the EKG Machines

Date: March 04, 2026

VA-26-00045513

PD Version Number:  1.1

# Contents

1.0	PRODUCT REQUIREMENTS	3

2.0	NOTICE OF THE FEDERAL ACCESSIBILITY LAW AFFECTING ALL INFORMATION AND COMMUNICATION TECHNOLOGY (ICT) PROCUREMENTS (SECTION 508)	3

2.1	SECTION 508 – INFORMATION AND COMMUNICATION TECHNOLOGY (ICT) STANDARDS	4

2.2	COMPATABILITY WITH ASSISTIVE TECHNOLOGY	4

2.3	ACCEPTANCE AND ACCEPTANCE TESTING	4

3.0	INFORMATION TECHNOLOGY USING ENERGY-EFFICIENT PRODUCTS	5

4.0	GENERAL REQUIREMENTS	5

4.1	VA TECHNICAL REFERENCE MODEL	5

4.2	SOCIAL SECURITY NUMBER (SSN) REDUCTION	5

4.3	INTERNET PROTOCOL VERSION 6 (IPV6)	6

# PRODUCT REQUIREMENTS

The El Paso Veterans Affairs Health Care System (EPVAHCS) is seeking comprehensive support services for its GE Muse systems. These systems are critical for managing and processing patient cardiovascular data, including ECG tests, which are essential for accurate diagnosis and effective treatment planning. The support services will encompass a range of functionalities, including the maintenance and optimization of the HL7 and DICOM interfaces, database management, data storage solutions, and application support.

Ensuring the operational integrity of our servers and computers is paramount, as any downtime could significantly impact patient care and the efficiency of our cardiology clinic. The GE Muse application plays a vital role in facilitating seamless communication and data exchange between various healthcare systems, thereby enhancing the quality of care provided to our veterans.

The EPVAHCS is specifically requesting dedicated service support for the GE server and workstation applications, as well as the HL7 and DICOM interfaces. This support will not only ensure that our systems remain fully operational but also enable us to adapt to evolving technological needs and maintain compliance with healthcare regulations. By investing in these support services, we aim to uphold our commitment to delivering timely and effective cardiovascular care to our patients, thereby improving health outcomes and overall patient satisfaction.

The contractor shall provide all tools, equipment, technical information, Technical Support for GE Muse software, servers, computers, Interface, applications, database, to include travel and personnel for proper assembly, repairs, and troubleshooting of systems and equipment.

Characteristics Of Service:

- GE Muse Server(s)

- GE Muse Web Client

- GE Muse Software

- DICOM or HL7 & Interface to Vista Imaging

The maintenance shall ensure that cardiology and biomedical engineering staff have access to email and phone technical support, as well as the latest updates to the GE Muse software. GE Muse is used as the Picture Archiving and Communication System (PACS) at the El Paso VA Medical Center cardiology department.

The following new software versions will be provided as they are released:

- GE Muse – Software used by the cardiology clinic by end users. Updates to software include:

  - Software and Quality Updates

  - Software Upgrades

  - Workflow Services

  - Installation Support

The technical support representative shall be qualified and have the required levels of professional and technical experience to perform the work required by this Purchase Description (PD). Email and phone technical support will be provided, enabling the cardiology and biomedical engineering staff to contact GE Healthcare for assistance with the following tasks:

- Software Upgrades and Updates with implementation and clinical education for every software upgrade.

- Phone Clinical Application Support: M-F, 8am-5pm

- Installation Support: Mon-Fri, 08:00AM-05:00PM

The Contractor is responsible for notifying the VA Program Manager, Contract Specialist and Contracting Officer in advance of arrival and departure for delivery and installation, providing review of all product and site conditions before signing off. Work is to be performed between 8:00am and 4:30pm during the week unless specifically coordinated otherwise, which can include weekends.

The contractor shall be responsible for any damage caused by delivery, installation, and removal while at the facility. The contractor shall ensure its processes do not create an unclean or unsafe environment for the EPVAHCS or damage any federal government equipment.

The period of performance will be a 12-month base period with four 12-month option periods.

# NOTICE OF THE FEDERAL ACCESSIBILITY LAW AFFECTING ALL INFORMATION AND COMMUNICATION TECHNOLOGY (ICT) PROCUREMENTS (SECTION 508)

On January 18, 2017, the Architectural and Transportation Barriers Compliance Board (Access Board) revised and updated, in a single rulemaking, standards for electronic and information technology developed, procured, maintained, or used by Federal agencies covered by Section 508 of the Rehabilitation Act of 1973, as well as our guidelines for telecommunications equipment and customer premises equipment covered by Section 255 of the Communications Act of 1934. The revisions and updates to the Section 508-based standards and Section 255-based guidelines are intended to ensure that information and communication technology (ICT) covered by the respective statutes is accessible to and usable by individuals with disabilities.

## SECTION 508 – INFORMATION AND COMMUNICATION TECHNOLOGY (ICT) STANDARDS

The Section 508 standards established by the Access Board are incorporated into, and made part of all VA orders, solicitations and purchase orders developed to procure ICT.  These standards are found in their entirety at: https://www.access-board.gov/guidelines-and-standards/communications-and-it/about-the-ict-refresh/final-rule/text-of-the-standards-and-guidelines.  A printed copy of the standards will be supplied upon request.

Federal agencies must comply with the updated Section 508 Standards beginning on January 18, 2018. The Final Rule as published in the Federal Register is available from the Access Board: https://www.access-board.gov/guidelines-and-standards/communications-and-it/about-the-ict-refresh/final-rule.

The Contractor shall comply with “508 Chapter 2: Scoping Requirements” for all electronic ICT and content delivered under this contract. Specifically, as appropriate for the technology and its functionality, the Contractor shall comply with the technical standards marked here:

E205 Electronic Content – (Accessibility Standard -WCAG 2.0 Level A and AA Guidelines)

E204 Functional Performance Criteria

E206 Hardware Requirements

E207 Software Requirements

E208 Support Documentation and Services Requirements

## COMPATIBILITY WITH ASSISTIVE TECHNOLOGY

The standards do not require installation of specific accessibility-related software or attachment of an assistive technology device. Section 508 requires that ICT be compatible with such software and devices so that ICT can be accessible to and usable by individuals using assistive technology, including but not limited to screen readers, screen magnifiers, and speech recognition software.

## ACCEPTANCE AND ACCEPTANCE TESTING

Deliverables resulting from this solicitation will be accepted based in part on satisfaction of the Section 508 Chapter 2: Scoping Requirements standards identified above.

The Government reserves the right to test for Section 508 Compliance before delivery.  The Contractor shall be able to demonstrate Section 508 Compliance upon delivery.

# INFORMATION TECHNOLOGY USING SUSTAINABLE PRODUCTS AND SERVICES

The Contractor shall comply with Sections 524 and Sections 525 of the Energy Independence and Security Act of 2007; Section 104 of the Energy Policy Act of 2005; Executive Order 13221, “Energy-Efficient Standby Power Devices,” dated August 2, 2001; and the Federal Acquisition Regulation (FAR) to provide ENERGY STAR®, Federal Energy Management Program (FEMP) designated, and products and services that meet EPA Recommendations of Specifications, Standards, and Ecolabels in effect as of October 2023 (e.g. Electronic Product Environmental Assessment Tool (EPEAT) registered products) in providing information technology products and/or services. The Green Procurement Compilation (GPC) website available at Green Procurement Compilation - GSA Sustainable Facilities Tool (sftool.gov) provides a comprehensive list of sustainable products and services and sustainable acquisition guidance.

The Contractor shall provide products that meet the definition of sustainable products and services if the products are delivered to the Government; furnished by the Contractor for use by the Government, incorporated into the construction of a public building or public work, or acquired by the Contractor for use in performing services under a Government contract where the cost of the products is a direct cost to a Government contract (versus costs which are normally applied to a contractor's general and administrative expenses or indirect costs).

Specifically, the Contractor shall provide sustainable products under this contract in accordance with FAR 52.223-23.  The applicable products with their associated certification, designation, and/or ecolabel (e.g. Energy Star, FEMP, and/or EPEAT) are identified within the product specifications in this PD.  The Contractor shall ensure proposed products meet the certification(s) / designation(s) / ecolabel(s) specified at the time of submission of proposals and at the time of award.

# SHIPMENT OF HARDWARE OR EQUIPMENT

Inspection:  Destination
Acceptance:  Destination
Free on Board (FOB):  Destination

Ship To and Mark For:

Primary POC:

Name: Peter J. Castillo

Address Name: El Paso VAHCS

Street Address: 5001 N. Piedras St. El Paso, TX 79930

Voice:	915-564-6100 Ext. 5555

Email:	peter.castillo1@va.gov

## SPECIAL SHIPPING INSTRUCTIONS

Prior to shipping, Contractor shall notify Site POCs, by phone followed by email, of all incoming deliveries including line-by-line details for review of requirements.  The Contractor shall not make any changes to the delivery schedule at the request of Site POC.

Contractors shall coordinate deliveries with Site POCs before shipment of <hardware> hardware to ensure sites have adequate storage space.

All shipments, either single or multiple container deliveries, shall bear the VA IFCAP Purchase Order number (if applicable) on external shipping labels and associated manifests or packing lists.  In the case of multiple container deliveries, a statement readable near the VA IFCAP PO or Contract number shall indicate total number of containers for the complete shipment (e.g. “Package 1 of 2”), clearly readable on manifests and external shipping labels.

# GENERAL REQUIREMENTS

## VA TECHNICAL REFERENCE MODEL

The Contractor shall comply with the VA OIT Technical Reference Model (VA TRM). Compliance with the VA TRM is achieved by using only technologies and standards that are listed as approved for use in the VA TRM.  The Contractor shall provide all necessary information requested by VA to ensure TRM approval is obtained prior to use on VA’s network.

## ZERO TRUST – VA CRITICAL SECURITY CONTROLS

VA has established minimum mandatory security requirements and requires that any network connected software system or service must meet the VA Critical Security Controls as outlined in the VA Memorandum, “VA Security Controls”, https://www.voa.va.gov/DocumentView.aspx?DocumentID=5010.  VA Critical Security Controls identify the minimum mandatory requirements that must be implemented across all VA enterprise infrastructure, cloud computing environments, information systems, networks, and specialized devices (medical devices/systems, special-purpose systems, and research scientific computing devices) that process, store, and/or transmit VA data.  Effective July 1, 2025, the Contractor shall implement these VA Critical Security Controls, within any network connected software system or service prior to being authorized for use in the VA.  This functional requirement is not negotiable, and Plan of Action & Milestones (POAM) will not be accepted in the event these controls cannot be implemented for new systems.  Critical Security Controls are intended to increase VA’s security posture and provide security and privacy risk visibility into the VA network and is not a new requirement.  The Contractor’s failure to maintain these VA Critical Controls after implementation will result in VA discontinuing the use of the system.

## SOCIAL SECURITY NUMBER (SSN) REDUCTION

The Contractor solution shall support the Social Security Number (SSN) Fraud Prevention Act (FPA) of 2017 which prohibits the inclusion of SSNs on any document sent by mail.  The Contractor support shall also be performed in accordance with Section 240 of the Consolidated Appropriations Act (CAA) 2018, enacted March 23, 2018, which mandates VA to discontinue using SSNs to identify individuals in all VA information systems as the Primary Identifier.  The Contractor shall ensure that any new IT solution discontinues the use of SSN as the Primary Identifier to replace the SSN with the Integrated Control Number (ICN) in all VA information systems for all individuals.  The Contractor shall ensure that all Contractor delivered applications and systems integrate with the VA Master Person Index (MPI) for identity traits to include the use of the ICN as the Primary Identifier. The Contractor solution may only use a Social Security Number to identify an individual in an information system if and only if the use of such number is required to obtain information VA requires from an information system that is not under the jurisdiction of VA.

## INTERNET PROTOCOL VERSION 6 (IPV6)

The Contractor solution shall support IPv6-Only based upon the memo issued by the Office of Management and Budget (OMB) on November 19, 2020 (https://www.whitehouse.gov/wp-content/uploads/2020/11/M-21-07.pdf). Which defines IPv6-only as the state of an operational system or service when IPv4 protocol functions (addressing, packet forwarding) are not in use. The NIST USGv6 profile defines technical requirements for a product to be capable of operating in IPv6-Only environments.  IPv6-Only technology, in accordance with the USGv6 Program (https://www.nist.gov/programs-projects/usgv6-program/usgv6-revision-1), NIST Special Publication (SP) 500-267B Revision 1 “USGv6 Profile” (https://doi.org/10.6028/NIST.SP.500-267Br1), and NIST SP 800-119 “Guidelines for the Secure Deployment of IPv6” (https://doi.org/10.6028/NIST.SP.800-119), compliance shall be included in all IT infrastructures, application designs, application development, operational systems and sub-systems, and their integration.  In addition to the above requirements, all devices, applications, and systems shall support all applicable functionality on native IPv6-Only as well as dual stack (IPv6 / IPv4) connectivity without additional memory or other resources being provided by the Government, so that they can function in a mixed environment.  All public/external facing servers and services (e.g., web, email, DNS, ISP services, etc.) shall support native IPv6-Only and dual stack (IPv6 / IPv4) users and all internal infrastructure and applications shall support using native IPv6-Only and dual stack (IPv6 / IPv4) for all functionality and operations.

## SOFTWARE AND LICENSING REQUIREMENTS

(If this acquisition is a software product and/or solution acquired as subscription service within an XaaS (Anything as a Service) environment (not via licenses) or this acquisition does not involve any aspect of a software solution or software product at all, this paragraph and Section can be removed entirely from the PD.  Any questions regarding this language should be addressed to Brandon Storms, Office of Strategic Sourcing, at Brandon.Storms@va.gov or via the VA OSS CECM All Staff (GAL Distribution Group.)

The Contractor shall be responsible for the provision of all software licenses and any associated licensing maintenance required for any development, delivery, integration, operation, and/or maintenance associated with its proposed application(s), software products, software solution, and/or system including, but not limited to, any and all application(s), software and/or software products that comprise, are a part of, or integrate with the Contractor’s proposed application(s), software products, software solution, and/or system for the life of any resulting contract.

## TRUSTED INTERNET CONNECTION (TIC)

The Contractor solution shall meet the requirements outlined in Office of Management and Budget Memorandum M-19-26, “Update to the Trusted Internet Connections (TIC) Initiative“ (https://www.whitehouse.gov/wp-content/uploads/2019/09/M-19-26.pdf), VA Directive 6513 “Secure External Connections”, and shall comply with the TIC 3.0 Core Guidance Documents, including all Volumes and TIC Use Cases, found at the Cybersecurity & Infrastructure Security Agency (CISA) (https://www.cisa.gov/publication/tic-30-core-guidance-documents.)

## POSITION/TASK RISK DESIGNATION LEVEL(S)

(For procurements which include continuous/full-time/on-site technical support being acquired with the hardware/software, additional security requirements may apply.  Refer to the Security Checklist, Appendix A in the 6500.6 Handbook, especially response to question 4 to determine if risk designation determinations are required.  If not, this section may be deleted.)

(All positions subject to investigation must be designated with a risk and sensitivity level based on the documented duties and responsibilities in accordance with VA Handbook 0710.  Using the results of the Position Designation Automated Tool (PDT) (also known as PDAT) performed for each applicable service being procured, complete the table below.  This “Service Task” Position Risk Designation correlates to individual Contractor personnel background investigation levels, depending upon which particular service task the contractor individual is working.

The PDT Tool is located at the following US Office of Personnel Management Website:  Position Designation Tool (opm.gov))

(If it is determined that position risk designations are required, include the below information, change the text to black and convert the text to “No Spacing” style.)

In accordance with VA Handbook 0710, Personnel Security and Suitability Program, the position sensitivity and the level of background investigation commensurate with the required level of access for the following tasks within the PWS are:

(List PD Item Numbers which contain the service(s) for this effort in the first column.  Click on the selection box in the appropriate Position Risk Designation column to mark an x to indicate the proper Risk Designation associated with each task based upon the PDT tool results.)

# Position Sensitivity and Background Investigation Requirements

# by Service Task/Item Number

| Service Task / Item Number | Tier1 / Low Risk | Tier 2 / Moderate Risk | Tier 4 / High Risk |
| --- | --- | --- | --- |
| 1.1 | ☐ | ☐ | ☐ |
| <1.2> | ☐ | ☐ | ☐ |
|  | ☐ | ☐ | ☐ |
|  | ☐ | ☐ | ☐ |

The Tasks identified above and the resulting Position Sensitivity and Background Investigation requirements identify, in effect, the Background Investigation requirements for Contractor individuals, based upon the tasks the particular Contractor individual will be working.  The submitted Contractor Staff Roster must indicate the required Background Investigation Level for each Contractor individual based upon the tasks the Contractor individual will be working, in accordance with their submitted proposal.

- POINTS OF CONTACT

VA Program Manager:

Name:	 Peter J. Castillo

Address: 5001 Piedras Street, El Paso, TX 79930

Voice:	915-564-6100 Ext. 5555

Email:	peter.castillo1@va.gov

Contracting Officer:

Name: Newton Maloba

Address: 23 Christopher Way, Eatontown, NJ 07701

Voice:	848-377-5305

Email:	newton.maloba@va.gov

Contract Specialist:

Name:	 Medin (Din) Boljevic

Address: 23 Christopher Way, Eatontown, NJ 07701

Voice:	848-377-5306

Email:	medin.boljevic@va.gov

(Maintenance/installation (warranty) contracts that may involve a contractor’s access or exposure to VA sensitive information, in particular to sensitive personal information while implementing their contract services, requires the inclusion of the requirements of Sections B1 through B5 in Addendum B accordingly)

(For procurements which include continuous/full-time/on-site technical support being acquired with the hardware/software, additional information security requirements may apply.  All Sections in this Addendum B must be reviewed to determine applicability to the requirements of the effort within the PD.  Any Sections (B1 through B14) which DO NOT apply should be removed from this Addendum B.  Requiring Activities should see the inclusion instructions prior to each Section and remove any full Sections that DO NOT apply (indicating “N/A” underneath the Section Heading(s)).  The language within these Sections MUST NOT be modified.)

(For PD’s that do not require these additional security requirements, Addendum B may be deleted.)

# ADDENDUM B – VA INFORMATION AND INFORMATION SYSTEM SECURITY/ PRIVACY LANGUAGE

NOTE: In the event of a conflict, VAAR Security Clauses take precedence over the language in this Addendum B.

APPLICABLE SECTIONS FROM: VA NOTICE 24-12, APRIL 22, 2024, UPDATE TO VA HANDBOOK 6500.6, CONTRACT SECURITY, APPENDIX C VA INFORMATION AND INFORMATION SYSTEM SECURITY/PRIVACY LANGUAGE, FOR INCLUSION INTO CONTRACTS, AS APPROPRIATE

(All Sections in this Addendum B must be reviewed to determine applicability to the requirements of the effort within the PWS.  Any Sections (B1 through B14) which DO NOT apply should be removed from this Addendum B.  Requiring Activities should see the inclusion instructions prior to each Section and remove any full Sections that DO NOT apply (indicating “N/A” underneath the Section Heading(s)).  The language within these Sections MUST NOT be modified, except for the removal of instructional text provided prior to each Section)

GENERAL

This entire section applies to all acquisitions requiring any Information Security and Privacy language. Contractors, contractor personnel, Subcontractors and subcontractor personnel will be subject to the same federal laws, regulations, standards, VA directives and handbooks, as VA personnel regarding information and information system security and privacy.

VA INFORMATION CUSTODIAL LANGUAGE

(This entire section applies to all acquisitions requiring any Information Security and Privacy language)

- The Government shall receive unlimited rights to data/intellectual property first produced and delivered in the performance of this contract or order (hereinafter “contract”) unless expressly stated otherwise in this contract. This includes all rights to source code and all documentation created in support thereof. The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General. The primary clause used to define computer software license (not data/intellectual property first produced under this Contractor or order) is FAR 52.227-19, Commercial Computer Software License.

- Information made available to the Contractor by VA for the performance or administration of this contract will be used only for the purposes specified in the service agreement, SOW, PWS, PD, and/or contract. The Contractor shall not use VA information in any other manner without prior written approval from a VA Contracting Officer (CO). The primary clause used to define Government and Contractor data rights is FAR 52.227-14 Rights in Data – General.

- VA information will not be co-mingled with any other data on the Contractor’s information systems or media storage systems. The Contractor shall ensure compliance with Federal and VA requirements related to data protection, data encryption, physical data segregation, logical data segregation, classification requirements and media sanitization.

- VA reserves the right to conduct scheduled or unscheduled audits, assessments, or investigations of Contractor Information Technology (IT) resources to ensure information security is compliant with Federal and VA requirements. The Contractor shall provide all necessary access to records (including electronic and documentary materials related to the contracts and subcontracts) and support (including access to Contractor and Subcontractor staff associated with the contract) to VA, VA's Office Inspector General (OIG), and/or Government Accountability Office (GAO) staff during periodic control assessments, audits, or investigations.

- The Contractor may only use VA information within the terms of the contract and applicable Federal law, regulations, and VA policies. If new Federal information security laws, regulations or VA policies become applicable after execution of the contract, the parties agree to negotiate contract modification and adjustment necessary to implement the new laws, regulations, and/or policies.

- The Contractor shall not make copies of VA information except as specifically authorized and necessary to perform the terms of the contract. If copies are made for restoration purposes, after the restoration is complete, the copies shall be destroyed in accordance with VA Directive 6500, VA Cybersecurity Program and VA Information Security Knowledge Service.

- If a Veterans Health Administration (VHA) contract is terminated for default or cause with a business associate, the related local Business Associate Agreement (BAA) shall also be terminated and actions taken in accordance with VHA Directive 1605.05, Business Associate Agreements. If there is an executed national BAA associated with the contract, VA will determine what actions are appropriate and notify the contractor.

- The Contractor shall store and transmit VA sensitive information in an encrypted form, using VA-approved encryption tools which are, at a minimum, Federal Information Processing Standards (FIPS) 140-2, Security Requirements for Cryptographic Modules (or its successor) validated and in conformance with VA Information Security Knowledge Service requirements. The Contractor shall transmit VA sensitive information using VA approved Transport Layer Security (TLS) configured with FIPS based cipher suites in conformance with National Institute of Standards and Technology (NIST) 800-52, Guidelines for the Selection, Configuration and Use of Transport Layer Security (TLS) Implementations.

- The Contractor’s firewall and web services security controls, as applicable, shall meet or exceed VA’s minimum requirements.

- Except for uses and disclosures of VA information authorized by this contract for performance of the contract, the Contractor may use and disclose VA information only in two situations: (i) in response to a qualifying order of a court of competent jurisdiction after notification to VA CO (ii) with written approval from the VA CO. The Contractor shall refer all requests for, demands for production of or inquiries about, VA information and information systems to the VA CO for response.

- Notwithstanding the provision above, the Contractor shall not release VA records protected by Title 38 U.S.C. § 5705, Confidentiality of medical quality-assurance records and/or Title 38 U.S.C. § 7332, Confidentiality of certain medical records pertaining to drug addiction, sickle cell anemia, alcoholism or alcohol abuse or infection with Human Immunodeficiency Virus (HIV). If the Contractor is in receipt of a court order or other requests for the above-mentioned information, the Contractor shall immediately refer such court order or other requests to the VA CO for response.

- Information made available to the Contractor by VA for the performance or administration of this contract or information developed by the Contractor in performance or administration of the contract will be protected and secured in accordance with VA Directive 6500 and Identity and Access Management (IAM) Security processes specified in the VA Information Security Knowledge Service.

- Any data destruction done on behalf of VA by a Contractor shall be done in accordance with National Archives and Records Administration (NARA) requirements as outlined in VA Directive 6300, Records and Information Management, VA Handbook 6300.1, Records Management Procedures, and applicable VA Records Control Schedules.

- The Contractor shall provide its plan for destruction of all VA data in its possession according to VA Directive 6500 and NIST 800-88, Guidelines for Media Sanitization prior to termination or completion of this contract. If directed by the COR/CO, the Contractor shall return all Federal Records to VA for disposition.

- Any media, such as paper, magnetic tape, magnetic disks, solid state devices or optical discs that is used to store, process, or access VA information that cannot be destroyed shall be returned to VA. The Contractor shall hold the appropriate material until otherwise directed by the Contracting Officer’s Representative (COR) or CO. Items shall be returned securely via VA-approved methods. VA sensitive information must be transmitted utilizing VA-approved encryption tools which are validated under FIPS 140-2 (or its successor) and NIST 800-52. If mailed, the Contractor shall send via a trackable method (USPS, UPS, FedEx, etc.) and immediately provide the COR/CO with the tracking information. Self-certification by the Contractor that the data destruction requirements above have been met shall be sent to the COR/CO within 30 business days of termination of the contract.

- All electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) used to store, process or access VA information will not be returned to the Contractor at the end of lease, loan, or trade-in. Exceptions to this paragraph will only be granted with the written approval of the VA CO.

ACCESS TO VA INFORMATION AND VA INFORMATION SYSTEMS

(This section applies when any person requires access to information made available to the Contractor by VA for the performance or administration of this contract or information developed by the Contractor in performance or administration of the contract)

- A Contractor/Subcontractor shall request logical (technical) or physical access to VA information and VA information systems for their employees and subcontractors only to the extent necessary to perform the services specified in the solicitation or contract. This includes indirect entities, both affiliate of Contractor/Subcontractor and agent of Contractor/Subcontractor.

- Contractors and subcontractors shall sign the VA Information Security Rule of Behavior (ROB) before access is provided to VA information and information systems (see Section 4, Training, below). The ROB contains the minimum user compliance requirements and does not supersede any policies of VA facilities or other agency components which provide higher levels of protection to VA’s information or information systems. Users who require privileged access shall complete the VA elevated privilege access request processes before privileged access is granted.

- All Contractors and Subcontractors working with VA information are subject to the same security investigative and clearance requirements as those of VA appointees or employees who have access to the same types of information. The level and process of background security investigations for contractors shall be in accordance with VA Directive and Handbook 0710, Personnel Suitability and Security Program. The Office of Human Resources and Administration/Operations, Security and Preparedness (HRA/OSP) is responsible for these policies and procedures. Contract personnel who require access to classified information or information systems shall have an appropriate security clearance. Verification of a Security Clearance shall be processed through the Special Security Officer located in HRA/OSP. Contractors shall conform to all requirements stated in the National Industrial Security Program Operating Manual (NISPOM).

- All Contractors and subcontractors shall comply with conditions specified in VAAR 852.204-71(d); Contractor operations required to be in United States. All Contractors and Subcontractors working with VA information must be permanently located within a jurisdiction subject to the law of the United States or its Territories to the maximum extent feasible. If services are proposed to be performed abroad the Contractor must state where all non-U.S. services are provided. The Contractor shall deliver to VA a detailed plan specifically addressing communications, personnel control, data protection and potential legal issues. The plan shall be approved by the COR/CO in writing prior to access being granted.

- The Contractor shall notify the COR/CO in writing immediately (no later than 24 hours) after personnel separation or occurrence of other causes. Causes may include the following:

  - Contractor/Subcontractor personnel no longer has a need for access to VA information or VA information systems.

  - Contractor/Subcontractor personnel are terminated, suspended, or otherwise has their work on a VA project discontinued for any reason.

  - Contractor believes their own personnel or Subcontractor personnel may pose a threat to their company’s working environment or to any company-owned property. This includes Contractor-owned assets, buildings, confidential data, customers, employees, networks, systems, trade secrets and/or VA data.

  - Any previously undisclosed changes to Contractor/Subcontractor background history are brought to light, including but not limited to changes to background investigation or employee record.

  - Contractor/Subcontractor personnel have their authorization to work in the United States revoked.

  - Agreement by which Contractor provides products and services to VA has either been fulfilled or terminated, such that VA can cut off electronic and/or physical access for Contractor personnel.

- In such cases of contract fulfillment, termination, or other causes; the Contractor shall take the necessary measures to immediately revoke access to VA network, property, information, and information systems (logical and physical) by Contractor/Subcontractor personnel. These measures include (but are not limited to): removing and then securing Personal Identity Verification (PIV) badges and PIV – Interoperable (PIV-I) access badges, VA-issued photo badges, credentials for VA facilities and devices, VA-issued laptops, and authentication tokens. Contractors shall notify the appropriate VA COR/CO immediately to initiate access removal.

- Contractors/Subcontractors who no longer require VA accesses will return VA-issued property to VA. This property includes (but is not limited to): documents, electronic equipment, keys, and parking passes. PIV and PIV-I access badges shall be returned to the nearest VA PIV Badge Issuance Office. Once they have had access to VA information, information systems, networks and VA property in their possessions removed, Contractors shall notify the appropriate VA COR/CO.

TRAINING

(This entire section applies to all acquisitions which include section 3)

- All contractors and subcontractors requiring access to VA information and VA information systems shall successfully complete the following before being granted access to VA information and its systems:

  - VA Privacy and Information Security Awareness and Rules of Behavior course (Talent Management System (TMS) #10176) initially and annually thereafter.

  - Sign and acknowledge (electronically through TMS #10176) understanding of and responsibilities for compliance with the Organizational Rules of Behavior, relating to access to VA information and information systems initially and annually thereafter; and

  - Successfully complete any additional cyber security or privacy training, as required for VA personnel with equivalent information system or information access [to be defined by the VA program official and provided to the VA CO for inclusion in the solicitation document – i.e., any role-based information security training].

- The Contractor shall provide to the COR/CO a copy of the training certificates and certification of signing the Organizational Rules of Behavior for each applicable employee within five days of the initiation of the contract and annually thereafter, as required.

- Failure to complete the mandatory annual training is grounds for suspension or termination of all physical or electronic access privileges and removal from work on the contract until such time as the required training is complete.

SECURITY INCIDENT INVESTIGATION

(This entire section applies to all acquisitions requiring any Information Security and Privacy language)

- The Contractor, Subcontractor, their employees, or business associates shall immediately (within one hour) report suspected security / privacy incidents to the VA OIT’s Enterprise Service Desk (ESD) by calling (855) 673-4357 (TTY: 711). The ESD is OIT’s 24/7/365 single point of contact for IT-related issues. After reporting to the ESD, the Contractor, Subcontractor, their employees, or business associates shall, within one hour, provide the COR/CO the incident number received from the ESD.

- To the extent known by the Contractor/Subcontractor, the Contractor/Subcontractor's notice to VA shall identify the information involved and the circumstances surrounding the incident, including the following:

  - The date and time (or approximation of) the Security Incident occurred.

  - The names of individuals involved (when applicable).

  - The physical and logical (if applicable) location of the incident.

  - Why the Security Incident took place (i.e., catalyst for the failure).

  - The amount of data belonging to VA believed to have been compromised.

  - The remediation measures the Contractor is taking to ensure no future incidents of a similar nature.

- After the Contractor has provided the initial detailed incident summary to VA, they will continue to provide written updates on any new and relevant circumstances or facts they discover. The Contractor, Subcontractor, and their employes shall fully cooperate with VA or third-party entity performing an independent risk analysis on behalf of VA. Failure to cooperate may be deemed a material breach and grounds for contract termination.

- VA IT Contractors shall follow VA Handbook 6500, Risk Management Framework for VA Information Systems VA Information Security Program, and VA Information Security Knowledge Service guidance for implementing an Incident Response Plan or integrating with an existing VA implementation.

- In instances of theft or break-in or other criminal activity, the Contractor/Subcontractor must concurrently report the incident to the appropriate law enforcement entity (or entities) of jurisdiction, including the VA OIG, and the VA Office of Security and Law Enforcement. The Contractor, its employees, and its Subcontractors and their employees shall cooperate with VA and any law enforcement authority responsible for the investigation and prosecution of any possible criminal law violation(s) associated with any incident. The Contractor/Subcontractor shall cooperate with VA in any civil litigation to recover VA information, obtain monetary or other compensation from a third party for damages arising from any incident, or obtain injunctive relief against any third party arising from, or related to, the incident.

- The Contractor shall comply with VA Handbook 6500.2, Management of Breaches Involving Sensitive Personal Information, which establishes the breach management policies and assigns responsibilities for the oversight, management and reporting procedures associated with managing of breaches.

- With respect to unsecured Protected Health Information (PHI), the Contractor is deemed to have discovered a data breach when the Contractor knew or should have known of breach of such information. When a business associate is part of VHA contract, notification to the covered entity (VHA) shall be made in accordance with the executed BAA.

- If the Contractor or any of its agents fails to protect VA sensitive personal information or otherwise engages in conduct which results in a data breach involving any VA sensitive personal information the Contractor/Subcontractor processes or maintains under the contract; the Contractor shall pay liquidated damages to the VA as set forth in clause 852.211-76, Liquidated Damages—Reimbursement for Data Breach Costs.

INFORMATION SYSTEM DESIGN AND DEVELOPMENT

(This entire section applies to information systems, systems, major applications, minor applications, enclaves, and platform information technologies (to include the subcomponents of each) designed or developed for or on behalf of VA by any non-VA entity)

- Information systems designed or developed on behalf of VA at non-VA facilities shall comply with all applicable Federal law, regulations, and VA policies. This includes standards for the protection of electronic Protected Health Information (PHI), outlined in 45 C.F.R. Part 164, Subpart C and information and system security categorization level designations in accordance with FIPS 199, Standards for Security Categorization of Federal Information and Information Systems and FIPS 200, Minimum Security Requirements for Federal Information Systems. Baseline security controls shall be implemented commensurate with the FIPS 199 system security categorization (reference VA Handbook 6500 and VA Trusted Internet Connections (TIC) Architecture).

- Contracted new developments require creation, testing, evaluation, and authorization in compliance with VA Assessment and Authorization (A&A) processes in VA Handbook 6500 and VA Information Security Knowledge Service to obtain an Authority to Operate (ATO). VA Directive 6517, Risk Management Framework for Cloud Computing Services, provides the security and privacy requirements for cloud environments.

- VA IT Contractors, Subcontractors and third-party service providers shall address and/or integrate applicable VA Handbook 6500, VA Handbook 6517, Risk Management Framework for Cloud Computing Services and Information Security Knowledge Service specifications in delivered IT systems/solutions, products and/or services. If systems/solutions, products and/or services do not directly match VA security requirements, the Contractor shall work though the COR/CO to identify the VA organization responsible for governance or resolution. Contractors shall comply with FAR 39.1, specifically the prohibitions referenced.

- The Contractor (including producers and resellers) shall comply with Office of Management and Budget (OMB) M-22-18 and M-23-16 when using third-party software on VA information systems or otherwise affecting the VA information. This includes new software purchases and software renewals for software developed or modified by major version change after the issuance date of M-22-18 (September 14, 2022). The term “software” includes firmware, operating systems, applications and application services (e.g., cloud-based software), as well as products containing software. The Contractor shall provide a self-attestation that secure software development practices are utilized as outlined by Executive Order (EO)14028 and NIST Guidance. A third-party assessment provided by either a certified Federal Risk and Authorization Management Program (FedRAMP) Third Party Assessor Organization (3PAO) or one approved by the agency will be acceptable in lieu of a software producer's self-attestation.

- The Contractor shall ensure all delivered applications, systems and information systems are compliant with Homeland Security Presidential Directive (HSPD) 12 and VA Identity and Access management (IAM) enterprise identity management requirements as set forth in OMB M-19-17, M-05-24, FIPS 201-3, Personal Identity Verification (PIV) of Federal Employees and Contractors (or its successor), M-21-31 and supporting NIST guidance. This applies to Commercial Off-The-Shelf (COTS) product(s) that the Contractor did not develop, all software configurations and all customizations.

- The Contractor shall ensure all Contractor delivered applications and systems provide user authentication services compliant with VA Handbook 6500, VA Information Security Knowledge Service, IAM enterprise requirements and NIST 800-63, Digital Identity Guidelines, for direct, assertion-based authentication and/or trust-based authentication, as determined by the design and integration patterns. Direct authentication at a minimum must include Public Key Infrastructure (PKI) based authentication supportive of PIV and/or Common Access Card (CAC), as determined by the business need and compliance with VA Information Security Knowledge Service specifications.

- The Contractor shall use VA authorized technical security baseline configurations and certify to the COR that applications are fully functional and operate correctly as intended on systems in compliance with VA baselines prior to acceptance or connection into an authorized VA computing environment. If the Defense Information Systems Agency (DISA) has created a Security Technical Implementation Guide (STIG) for the technology, the Contractor may configure to comply with that STIG. If VA determines a new or updated VA configuration baseline needs to be created, the Contractor shall provide required technical support to develop the configuration settings. FAR 39.1 requires the population of operating systems and applications includes all listed on the NIST National Checklist Program Checklist Repository,

- The standard installation, operation, maintenance, updating and patching of software shall not alter the configuration settings from VA approved baseline configuration. Software developed for VA must be compatible with VA enterprise installer services and install to the default “program files” directory with silently install and uninstall. The Contractor shall perform testing of all updates and patching prior to implementation on VA systems.

- Applications designed for normal end users will run in the standard user context without elevated system administration privileges.

- The Contractor-delivered solutions shall reside on VA approved operating systems. Exceptions to this will only be granted with the written approval of the COR/CO.

- The Contractor shall design, develop, and implement security and privacy controls in accordance with the provisions of VA security system development life cycle outlined in NIST 800-37, Risk Management Framework for Information Systems and Organizations: A System Life Cycle Approach for Security and Privacy, VA Directive and Handbook 6500, and VA Handbook 6517.

- The Contractor shall comply with the Privacy Act of1974 (the Act), FAR 52.224-2 Privacy Act, and VA rules and regulations issued under the Act in the design, development, or operation of any system of records on individuals to accomplish a VA function.

- The Contractor shall ensure the security of all procured or developed information systems, systems, major applications, minor applications, enclaves and platform information technologies, including their subcomponents (hereinafter referred to as “Information Systems”) throughout the life of this contract and any extension, warranty, or maintenance periods. This includes security configurations, workarounds, patches, hotfixes, upgrades, replacements and any physical components which may be necessary to remediate all security vulnerabilities published or known to the Contractor anywhere in the information systems (including systems, operating systems, products, hardware, software, applications and firmware). The Contractor shall ensure security fixes do not negatively impact the Information Systems.

- When the Contractor is responsible for operations or maintenance of the systems, the Contractor shall apply the security fixes within the timeframe specified by the associated controls on the VA Information Security Knowledge Service. When security fixes involve installing third party patches (such as Microsoft OS patches or Adobe Acrobat), the Contractor shall provide written notice to the VA COR/CO that the patch has been validated as to not affecting the Systems within 10 business days.

INFORMATION SYSTEM HOSTING, OPERATION, MAINTENANCE OR USE

(This entire section applies to information systems, systems, major applications, minor applications, enclaves, and platform information technologies (cloud and non-cloud) hosted, operated, maintained, or used on behalf of VA at non-VA facilities)

- The Contractor shall comply with all Federal laws, regulations, and VA policies for Information systems (cloud and non-cloud) that are hosted, operated, maintained, or used on behalf of VA at non-VA facilities. Security controls for collecting, processing, transmitting, and storing of VA sensitive information, must be in place. The controls will be tested by VA or a VA sanctioned 3PAO and approved by VA prior to hosting, operation, maintenance or use of the information system or systems by or on behalf of VA. This includes conducting compliance risk assessments, security architecture analysis, routine vulnerability scanning, system patching, change management procedures and the completion of an acceptable contingency plan for each system. The Contractor’s security control procedures shall be the same as procedures used to secure VA-operated information systems.

- Outsourcing (Contractor facility, equipment, or staff) of systems or network operations, telecommunications services or other managed services require Assessment and Authorization (A&A) of the Contractor’s systems in accordance with VA Handbook 6500 as specified in VA Information Security Knowledge Service. Major changes to the A&A package may require reviewing and updating all the documentation associated with the change. The Contractor’s cloud computing systems shall comply with FedRAMP and VA Directive 6517 requirements.

- The Contractor shall return all electronic storage media (hard drives, optical disks, CDs, back-up tapes, etc.) on non-VA leased or non-VA owned IT equipment used to store, process or access VA information to VA in accordance with A&A package requirements. This applies when the contract is terminated or completed and prior to disposal of media. The Contractor shall provide its plan for destruction of all VA data in its possession according to VA Information Security Knowledge Service requirements and NIST 800-88. The Contractor shall send a self-certification that the data destruction requirements above have been met to the COR/CO within 30 business days of termination of the contract.

- All external internet connections to VA network involving VA information must be in accordance with VA Trusted Internet Connection (TIC) Reference Architecture and VA Directive and Handbook 6513, Secure External Connections and reviewed and approved by VA prior to implementation. Government-owned Contractor-operated systems, third party or business partner networks require a Memorandum of Understanding (MOU) and Interconnection Security Agreements (ISA).

- Contractor procedures shall be subject to periodic, announced, or unannounced assessments by VA officials, the OIG or a 3PAO. The physical security aspects associated with Contractor activities are also subject to such assessments. The Contractor shall report, in writing, any deficiencies noted during the above assessment to the VA COR/CO. The Contractor shall use VA’s defined processes to document planned remedial actions that address identified deficiencies in information security policies, procedures, and practices. The Contractor shall correct security deficiencies within the timeframes specified in the VA Information Security Knowledge Service.

- All major information system changes which occur in the production environment shall be reviewed by the VA to determine the impact on privacy and security of the system. Based on the review results, updates to the Authority to Operate (ATO) documentation and parameters may be required to remain in compliance with VA Handbook 6500 and VA Information Security Knowledge Service requirements.

- The Contractor shall conduct an annual privacy and security self-assessment on all information systems and outsourced services as required. Copies of the assessment shall be provided to the COR/CO. The VA/Government reserves the right to conduct assessment using government personnel or a third-party if deemed necessary. The Contractor shall correct or mitigate any weaknesses discovered during the assessment.

- VA prohibits the installation and use of personally owned or Contractor-owned equipment or software on VA information systems. If non-VA owned equipment must be used to fulfill the requirements of a contract, it must be stated in the service agreement, SOW, PWS, PD or contract. All security controls required for government furnished equipment must be utilized in VA approved Other Equipment (OE). Configuration changes to the Contractor OE, must be funded by the owner of the equipment. All remote systems must use a VA-approved antivirus software and a personal (host-based or enclave based) firewall with a VA-approved configuration. The Contractor shall ensure software on OE is kept current with all critical updates and patches. Owners of approved OE are responsible for providing and maintaining the anti-virus software and the firewall on the non-VA owned OE. Approved Contractor OE will be subject to technical inspection at any time.

- The Contractor shall notify the COR/CO within one hour of disclosure or successful exploits of any vulnerability which can compromise the confidentiality, integrity, or availability of the information systems. The system or effected component(s) need(s) to be isolated from the network. A forensic analysis needs to be conducted jointly with VA. Such issues will be remediated as quickly as practicable, but in no event longer than the timeframe specified by VA Information Security Knowledge Service. If sensitive personal information is compromised reference VA Handbook 6500.2 and Section 5, Security Incident Investigation.

- For cases wherein the Contractor discovers material defects or vulnerabilities impacting products and services they provide to VA, the Contractor shall develop and implement policies and procedures for disclosure to VA, as well as remediation. The Contractor shall, within 30 business days of discovery, document a summary of these vulnerabilities or defects. The documentation will include a description of the potential impact of each vulnerability and material defect, compensating security controls, mitigations, recommended corrective actions, root cause analysis and/or workarounds (i.e., monitoring). Should there exist any backdoors in the products or services they provide to VA (referring to methods for bypassing computer authentication), the Contractor shall provide the VA CO/CO written assurance they have permanently remediated these backdoors.

- All other vulnerabilities, including those discovered through routine scans or other assessments, will be remediated based on risk, in accordance with the remediation timelines specified by the VA Information Security Knowledge Service and/or the applicable timeframe mandated by Cybersecurity & Infrastructure Security Agency (CISA) Binding Operational Directive (BOD) 22-01 and BOD 19-02 for Internet-accessible systems. Exceptions to this paragraph will only be granted with the approval of the COR/CO.

SECURITY AND PRIVACY CONTROLS COMPLIANCE TESTING, ASSESSMENT AND AUDITING

(This entire section applies whenever section 6 or 7 is included)

- Should VA request it, the Contractor shall provide a copy of their (corporation’s, sole proprietorship’s, partnership’s, limited liability company (LLC), or other business structure entity’s) policies, procedures, evidence and independent report summaries related to specified cybersecurity frameworks (International Organization for Standardization (ISO), NIST Cybersecurity Framework (CSF), etc.). VA or its third-party/partner designee (if applicable) are further entitled to perform their own audits and security/penetration tests of the Contractor’s IT or systems and controls, to ascertain whether the Contractor is complying with the information security, network or system requirements mandated in the agreement between VA and the Contractor.

- Any audits or tests of the Contractor or third-party designees/partner VA elects to carry out will commence within 30 business days of VA notification. Such audits, tests and assessments may include the following: (a): security/penetration tests which both sides agree will not unduly impact Contractor operations; (b): interviews with pertinent stakeholders and practitioners; (c): document review; and (d): technical inspections of networks and systems the Contractor uses to destroy, maintain, receive, retain, or use VA information.

- As part of these audits, tests and assessments, the Contractor shall provide all information requested by VA. This information includes, but is not limited to, the following: equipment lists, network or infrastructure diagrams, relevant policy documents, system logs or details on information systems accessing, transporting, or processing VA data.

- The Contractor and at its own expense, shall comply with any recommendations resulting from VA audits, inspections and tests. VA further retains the right to view any related security reports the Contractor has generated as part of its own security assessment. The Contractor shall also notify VA of the existence of any such security reports or other related assessments, upon completion and validation.

- VA appointed auditors or other government agency partners may be granted access to such documentation on a need-to-know basis and coordinated through the COR/CO. The Contractor shall comply with recommendations which result from these regulatory assessments on the part of VA regulators and associated government agency partners.

PRODUCT INTEGRITY, AUTHENTICITY, PROVENANCE, ANTI-COUNTERFEIT AND ANTI-TAMPERING

(This entire section applies when the acquisition involves any product (application, hardware, or software) or when section 6 or 7 is included)

- The Contractor shall comply with Code of Federal Regulations (CFR) Title 15 Part 7, “Securing the Information and Communications Technology and Services (ICTS) Supply Chain”, which prohibits ICTS Transactions from foreign adversaries. ICTS Transactions are defined as any acquisition, importation, transfer, installation, dealing in or use of any information and communications technology or service, including ongoing activities, such as managed services, data transmission, software updates, repairs or the platforming or data hosting of applications for consumer download.

- When contracting terms require the Contractor to procure equipment, the Contractor shall purchase or acquire the equipment from an Original Equipment Manufacturer (OEM) or an authorized reseller of the OEM. The Contractor shall attest that equipment procured from an OEM or authorized reseller or distributor are authentic. If procurement is unavailable from an OEM or authorized reseller, the Contractor shall submit in writing, details of the circumstances prohibiting this from happening and procure a product waiver from the VA COR/CO.

- All Contractors shall establish, implement, and provide documentation for risk management practices for supply chain delivery of hardware, software (to include patches) and firmware provided under this agreement. Documentation will include chain of custody practices, inventory management program, information protection practices, integrity management program for sub-supplier provided components, and replacement parts requests. The Contractor shall make spare parts available. All Contractor(s) shall specify how digital delivery for procured products, including patches, will be validated and monitored to ensure consistent delivery. The Contractor shall apply encryption technology to protect procured products throughout the delivery process.

- If a Contractor provides software or patches to VA, the Contractor shall publish or provide a hash conforming to the FIPS Security Requirements for Cryptographic Modules (FIPS 140-2 or successor).

- The Contractor shall provide a software bill of materials (SBOM) for procured (to include licensed products) and consist of a list of components and associated metadata which make up the product. SBOMs must be generated in one of the data formats defined in the National Telecommunications and Information Administration (NTIA) report “The Minimum Elements for a Software Bill of Materials (SBOM).”

- Contractors shall use or arrange for the use of trusted channels to ship procured products, such as U.S. registered mail and/or tamper-evident packaging for physical deliveries.

- Throughout the delivery process, the Contractor shall demonstrate a capability for detecting unauthorized access (tampering).

- The Contractor shall demonstrate chain-of-custody documentation for procured products and require tamper-evident packaging for the delivery of this hardware.

VIRUSES, FIRMWARE, AND MALWARE

(This entire section applies when the acquisition involves any product (application, hardware, or software) or when section 6 or 7 is included)

- The Contractor shall execute due diligence to ensure all provided software and patches, including third-party patches, are free of viruses and/or malware before releasing them to or installing them on VA information systems.

- The Contractor warrants it has no knowledge of and did not insert, any malicious virus and/or malware code into any software or patches provided to VA which could potentially harm or disrupt VA information systems. The Contractor shall use due diligence, if supplying third-party software or patches, to ensure the third-party has not inserted any malicious code and/or virus which could damage or disrupt VA information systems.

- The Contractor shall provide or arrange for the provision of technical justification as to why any “false positive” hit has taken place to ensure their code’s supply chain has not been compromised. Justification may be required, but is not limited to, when install files, scripts, firmware, or other Contractor-delivered software solutions (including third-party install files, scripts, firmware, or other software) are flagged as malicious, infected, or suspicious by an anti-virus vendor.

- The Contractor shall not upload (intentionally or negligently) any virus, worm, malware or any harmful or malicious content, component and/or corrupted data/source code (hereinafter “virus or other malware”) onto VA computer and information systems and/or networks. If introduced (and this clause is violated), upon written request from the VA CO, the Contractor shall:

  - Take all necessary action to correct the incident, to include any and all assistance to VA to eliminate the virus or other malware throughout VA’s information networks, computer systems and information systems; and

  - Use commercially reasonable efforts to restore operational efficiency and remediate damages due to data loss or data integrity damage, if the virus or other malware causes a loss of operational efficiency, data loss, or damage to data integrity.

CRYPTOGRAPHIC REQUIREMENT

(This entire section applies whenever the acquisition includes section 6 or 7 is included)

- The Contractor shall document how the cryptographic system supporting the Contractor’s products and/or services protect the confidentiality, data integrity, authentication and non-repudiation of devices and data flows in the underlying system.

- The Contractor shall use only approved cryptographic methods as defined in FIPS 140-2 (or its successor) and NIST 800-52 standards when enabling encryption on its products.

- The Contractor shall provide or arrange for the provision of an automated remote key-establishment method which protects the confidentiality and integrity of the cryptographic keys.

- The Contractor shall ensure emergency re-keying of all devices can be remotely performed within 30 business days.

- The Contractor shall provide or arrange for the provision of a method for updating cryptographic primitives or algorithms.

PATCHING GOVERNANCE

(This entire section applies whenever the acquisition includes section 7 is included)

- The Contractor shall provide documentation detailing the patch management, vulnerability management, mitigation and update processes (to include third-party) prior to the connection of electronic devices, assets or equipment to VA’s assets. This documentation will include information regarding the following:

  - The resources and technical capabilities to sustain the program or process (e.g., how the integrity of a patch is validated by VA); and.

  - The approach and capability to remediate newly reported zero-day vulnerabilities for Contractor products.

- The Contractor shall verify and provide documentation all procured products (including third-party applications, hardware, software, operating systems, and firmware) have appropriate updates and patches installed prior to delivery to VA.

- The Contractor shall provide or arrange the provision of appropriate software and firmware updates to remediate newly discovered vulnerabilities or weaknesses for their products and services within 30 days of discovery. Updates to remediate critical or emergent vulnerabilities will be provided within seven business days of discovery. If updates cannot be made available by Contractor within these time periods, the Contractor shall submit mitigations, methods of exploit detection and/or workarounds to the COR/CO prior to the above deadlines.

- The Contractor shall provide or arrange for the provision of appropriate hardware, software and/or firmware updates, when those products, including open-source software, are provided to the VA, to remediate newly discovered vulnerabilities or weaknesses. Remediations of products or services provided to the VA’s system environment must be provided within 30 business days of availability from the original supplier and/or patching source. Updates to remediate critical vulnerabilities applicable to the Contractor’s use of the third-party product in its system environment will be provided within seven business days of availability from the original supplier and/or patching source. If applicable third-party updates cannot be integrated, tested and made available by Contractor within these time periods, mitigations and/or workarounds will be provided to the COR/CO before the above deadlines.

SPECIALIZED DEVICES/SYSTEMS (MEDICAL DEVICES, SPECIAL PURPOSE SYSTEMS, RESEARCH SCIENTIFIC COMPUTING)

(This entire section applies when the acquisition includes one or more Medical Device, Special Purpose System or Research Scientific Computing Device. If appropriate, ensure selected clauses from section 6 or 7 and 8 through 12 are included)

- Contractor supplies/delivered Medical Devices, Special Purpose Systems-Operational Technology (SPS-OT) and Research Scientific Computing Devices shall comply with all applicable Federal law, regulations, and VA policies. New developments require creation, testing, evaluation, and authorization in compliance with processes specified on the Specialized Device Cybersecurity Department Enterprise Risk Management (SDCD-ERM) Portal, VA Directive 6550, Pre-Procurement Assessment and Implementation of Medical Devices/Systems, VA Handbook 6500, and the VA Information Security Knowledge Service. Deviations from Federal law, regulations, and VA Policy are identified and documented as part of VA Directive 6550 and/or the VA Enterprise Risk Analysis (ERA) processes for Specialized Devices/Systems processes.

- All Contractors and third-party service providers shall address and/or integrate applicable VA Handbook 6500 and Information Security Knowledge Service specifications in delivered IT systems/solutions, products and/or services. If systems/solutions, products and/or services do not directly match VA security requirements, the Contractor shall work though the COR/CO for governance or resolution.

- The Contractor shall certify to the COR/CO that devices/systems that have completed the VA Enterprise Risk Analysis (ERA) process for Specialized Devices/Systems are fully functional and operate correctly as intended. Devices/systems must follow the VA ERA authorized configuration prior to acquisition and connection to the VA computing environment. If VA determines a new VA ERA needs to be created, the Contractor shall provide required technical support to develop the configuration settings. Major changes to a previously approved device/system will require a new ERA.

- The Contractor shall comply with all practices documented by the Food Drug and Administration (FDA) Premarket Submission for Management of Cybersecurity in Medical Devices and Postmarket Management of Cybersecurity in Medical Devices.

- The Contractor shall design devices capable of accepting all applicable security patches with or without the support of the Contractor personnel. If patching can only be completed by the Contractor, the Contractor shall commit the resources needed to patch all applicable devices at all VA locations. If unique patching instructions or packaging is needed, the Contractor shall provide the necessary information in conjunction with the validation/testing of the patch. The Contractor shall apply security patches within 30 business days of the patch release and have a formal tracking process for any security patches not implemented to include explanation when a device cannot be patched.

- The Contractor shall provide devices able to install and maintain VA-approved antivirus capabilities with the capability to quarantine files and be updated as needed in response to incidents. Alternatively, a VA-approved whitelisting application may be used when the Contractor cannot install an anti-virus / anti-malware application.

- The Contractor shall verify and document all software embedded within the device does not contain any known viruses or malware before delivery to or installation at a VA location.

- Devices and other equipment or systems containing media (hard drives, optical disks, solid state, and storage via chips/firmware) with VA sensitive information will be returned to the Contractor with media removed. When the contract requires return of equipment, the options available to the Contractor are the following:

  - The Contractor shall accept the system without the drive, firmware and solid state.

  - VA’s initial device purchase includes a spare drive or other replacement media which must be installed in place of the original drive at time of turn-in; or

  - Due to the highly specialized and sometimes proprietary hardware and software associated with the device, if it is not possible for VA to retain the hard drive, firmware, and solid state, then:

    - The equipment Contractor shall have an existing BAA if the device being traded in has sensitive information stored on it and hard drive(s) from the system are being returned physically intact.

    - Any fixed hard drive, Complementary Metal-Oxide-Semiconductor (CMOS), Programmable Read-Only Memory (PROM), solid state and firmware on the device must be non-destructively sanitized to the greatest extent possible without negatively impacting system operation. Selective clearing down to patient data folder level is recommended using VA approved and validated overwriting technologies/methods/tools. Applicable media sanitization specifications need to be pre-approved and described in the solicitation, contract, or order.

DATA CENTER PROVISIONS

(This entire section applies whenever the acquisition requires an interconnection to/from the VA network to/from a non-VA location)

- The Contractor shall ensure the VA network is accessed by in accordance with VA Directive 6500 and IAM security processes specified in the VA Information Security Knowledge Service.

- The Contractor shall ensure network infrastructure and data availability in accordance with VA information system business continuity procedures specified in the VA Information Security Knowledge Service.

- The Contractor shall ensure any connections to the internet or other external networks for information systems occur through managed interfaces utilizing VA approved boundary protection devices (e.g., internet proxies, gateways, routers, firewalls, guards or encrypted tunnels).

- The Contractor shall encrypt all traffic across the segment of the Wide Area Network (WAN) it manages and no unencrypted Out of Band (OOB) Internet Protocol (IP) traffic will traverse the network.

- The Contractor shall ensure tunnel endpoints are routable addresses at each VA operating site.

- The Contractor shall secure access from Local Area Networks (LANs) at co-located sites in accordance with VA TIC Reference Architecture, VA Directive and Handbook 6513, and MOU/ISA process specified in the VA Information Security Knowledge Service.