| SUBJECT* | WATER PURIFICATION SYSTEM HAMPTON VAMC/DIALYSIS |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 23667 |
| --- | --- |
| SOLICITATION NUMBER* | 36C24626Q0335 |
| RESPONSE DATE/TIME/ZONE | 02-17-2026 10am EASTERN TIME, NEW YORK, USA |
| ARCHIVE | 30 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | 4610 |
| NAICS CODE* | 333310 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Network Contracting Office 6 100 Emancipation Drive Hampton VA 23667 |
| POINT OF CONTACT* |  |

PLACE OF PERFORMANCE

| ADDRESS |  |
| --- | --- |
| POSTAL CODE |  |
| COUNTRY |  |

ADDITIONAL INFORMATION

| AGENCY’S URL |  |
| --- | --- |
| URL DESCRIPTION |  |
| AGENCY CONTACT’S EMAIL ADDRESS |  |
| EMAIL DESCRIPTION |  |

DESCRIPTION

VAMC Hampton is looking for a Water Purification system with the following:

-	Salient Characteristics

-	Water Purification System that can produce not less than 200 liters of water per hour, fitted with deionization capabilities

-	MODEL MEDICA-R 200

-	Make-up rate 15C 200 l/hr

-	Delivery flow rate – maximum115V, 60Hz 21 l/min @ 45psi (3bar) Inorganics (resistivity @25°C) >10 MΩ-cm (1)

-	Organics (TOC) - typical <30ppb -500 PPB (1) Bacteria - typical <10 CFU/ml (2)

-	Particles 0.2 μm filter Silica <0.05 mg/l

-	Notes: 1 with Ion-exchange cylinder installed (Nuclear or Hypex grade resin), 2 subject to correct operating and maintenance procedures

-	Feedwater Requirement:

-	Source: Tap water as detailed below Conductivity: <1000 μS/cm

-	Contaminant:

-	Measure:	Range:	Pretreatment:

-	Calcium: Ca ppm as CaCO3	<250	None3

-	>250	Softener or use very low RO recovery3

-	Free chlorine: Cl2ppm	<0.1	None

-	0.1 – 0.5	20 inch carbon block

-	>0.5	Cylinder of carbon sized correctly to obtain <0.1 ppm

-	Chloramine: Cl2 ppm	<0.1	None

-	0.1 – 0.5	20 inch high efficiency carbon (5)

-	>0.5	Cylinder of carbon sized correctly to obtain <0.1 ppm

-	Silica: SiO2 ppm	<30	None

-	>30	20 inch cartridge depth filter or use very low RO recovery (4)

-	Fouling index: FI	<10	None

-	10 to 20	20 inch cartridge depth filter

-	>20	Back washable media filter with a minimum flow rate of 20 l/min

-	Iron/manganese: Fe/Mn ppm	<0.05	None

-	>0.05	20 inch cartridge depth filter (5)

-	>0.1	Back-washable Fe filter (6)

-	Organics: TOC ppm C	<2	None

-	2 to 3	20 inch carbon block (4)

-	>3	Carbon sized Cylinder correctly for TOC demand (4) Temperature: 4 - 40°C (Recommended 15 - 25°C)

--	Flowrate: (requirement at 15°C) 20 l/min

-	Drain requirements: (gravity fall with air gap) 45 l/min

-	Feedwater pressure: 60psi (4bar) maximum, 30psi (2bar) minimum

-	Notes: 3) Check LSI, increase frequency of acid cleaning, 4) Increase frequency of alkaline cleaning,

-	5) Please use ELGA Part Numbers LA683 (Filter Housing) and LC172 (Carbon Filter), 6) Increase frequency of acid cleaning Wall mount canister filter will need (Wall backing required)

-	Must include Installation, Commissioning, Training

-	Electrical Requirements Mains Input: 115V, 60Hz

-	System control voltage (not including pumps and UV) 24V dc Power consumption (peak demand): 2000VA

-	Electrical protection rating 20 amps Noise level during recirculation <70dBA

-	***Note: System requires approximately 4' x 6' of floor space for tanks, filtration Unit, and filters.

-	-The floor drain must be a minimum of 3" and meet all applicable building codes and system requirements

-	-Actual Unit measurement 32" Wide x35"Deep x72"High