| SUBJECT* | Multicolor Scan Laser Photocoagulator |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 17042 |
| --- | --- |
| SOLICITATION NUMBER* | 36C24426Q0537 |
| RESPONSE DATE/TIME/ZONE | 05-05-2026 4:00 PM EASTERN TIME, NEW YORK, USA |
| ARCHIVE | 10 DAYS AFTER THE RESPONSE DATE |
| SET-ASIDE | SDVOSBC |  |
| PRODUCT SERVICE CODE* | 6540 |  |
| NAICS CODE* | 334510 |  |
| PLACE OF PERFORMANCE | Lebanon VA Medical Center |  |
|  | 1700 South Lincoln Ave. |  |
|  | Lebanon PA |  |
| POSTAL CODE | 17042 |  |
| COUNTRY | USA |  |

CONTACT INFORMATION

| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Network Contracting Office 4 1010 Delafield Road Pittsburgh, PA 15240 |
| --- | --- |
| POINT OF CONTACT* | Contracting Officer Andrew Taylor andrew.taylor3@va.gov 724-679-2327 |

ADDITIONAL INFORMATION

| AGENCY’S URL | https://www.va.gov/lebanon-health-care/ |
| --- | --- |
| URL DESCRIPTION | Lebanon VA Medical Center |
| AGENCY CONTACT’S EMAIL ADDRESS | andrew.taylor3@va.gov |
| EMAIL DESCRIPTION | Contracting Officer |

DESCRIPTION

Statement of Need

Multicolor Scan Laser Photocoagulator

Background:

The Lebanon VA Medical Center (LVAMC) has a requirement for one (1) Multicolor Scan Laser Photocoagulator and its accessories.

Scope of Work: The multicolor scan laser photocoagulator provides a variety of laser treatments including pan retinal photocoagulation for diabetic retinopathy and laser iridotomy for glaucoma with a scan slit lamp delivery unit.

Required Features:

- 532 nm (green) wavelength for treating retinal pathology

- 577 nm (yellow) laser

- 647 nm (red) wavelength this wavelength is used for photocoagulation of deep choroidal pathology

- Allows the selection of one, two, or three wavelengths, among green, yellow, and red. Will allow selection of the necessary color or combination of colors.

- 22 preprogrammed scan patterns to allow treatment of varying retinal pathologies.

- Will allow automated positioning of the scan pattern to the next region to undergo photocoagulation.

- Will include attachable models that incorporate SOLIC (Safety Impacts with Low Impact on Cornea).

- The scan spot size will continuously be variable from 100 to 500 µm 50 to 500 µm in single mode.

- Must provide momentary increase followed by a plateau and an immediate decrease to provide rapid and high-power laser emission on the scan patterns.

Main Body Specifications:

- Laser type: Solid state laser, Diode laser

- Wavelength:

- Green: 532 nm

- Yellow: 577 nm

- Red: 647 nm

- Power output: Green: 50 to 1,700 mW•1 Yellow: 50 to 1,500 mW Red: 50 to 800 mW•2

- Output type: Continuous wave

- Emission time: 0.01 to 1.00 s, 2.00 s, 3.00 s*3

- Interval time: 0.05 to 1.00 s*4

- Aiming beam: Red diode, 670 nm, max. 0.6 mW

- Cooling system: Air cooling

- Power supply: 115/230 V AC, 50/60 Hz

- Power consumption: 400VA

- Dimensions/Mass: 300 (VI/) x 480 (D) x 670 (H) mm / 35 kg 11.8 (VI/) x 18.9 (D) x 26.4 (H)" /77.1 lbs

Scan I Single Delivery Unit Specifications

- Spot size: 50 to 1,000 µm (slit lamp & attachable deliveries)

- Emission pattern: Single

- Type: Slit lamp delivery unit

- Attachable delivery unit

- Dimensions/Mass: slit lamp delivery with table

Delivery:

Delivery shall be within thirty (30) days from the time of award.  The Multicolor Scan Laser Photocoagulator shall be delivered to the Lebanon VAMC located at 1700 S. Lincoln Ave., Lebanon, PA 17042; Attn; Warehouse, Biomed will conduct their inspection before delivery to the Optometry Clinic.

End of Document