| SUBJECT* | 629-26-3-106-0321 Thyroid Equipment (VA-26-00059957) 36C256-26-AP-3049 POC: Tiffany Conner |
| --- | --- |

GENERAL INFORMATION

| CONTRACTING OFFICE’S ZIP CODE* | 71101 |
| --- | --- |
| SOLICITATION NUMBER* | 36C25626Q0638 |
| RESPONSE DATE/TIME/ZONE | 04-27-2026 10:00 CENTRAL TIME, CHICAGO, USA |
| ARCHIVE | 10 DAYS AFTER THE RESPONSE DATE |
| RECOVERY ACT FUNDS | N |
| SET-ASIDE |  |
| PRODUCT SERVICE CODE* | 6525 |
| NAICS CODE* | 334516 |
| CONTRACTING OFFICE ADDRESS | Department of Veterans Affairs Overton Brooks VA Medical Center (90C) 510 East Stoner Avenue Shreveport LA 71101 |
| POINT OF CONTACT* | CO Tiffany Conner tiffany.conner@va.gov (318) 990-4064 |

PLACE OF PERFORMANCE

| ADDRESS |  |
| --- | --- |
|  | Southeast Louisiana VHCS |
|  | 2400 Canal St. |
|  | New Orleans LA |
| POSTAL CODE | 70119 |
| COUNTRY | US |

ADDITIONAL INFORMATION

| AGENCY’S URL |  |
| --- | --- |
| URL DESCRIPTION |  |
| AGENCY CONTACT’S EMAIL ADDRESS |  |
| EMAIL DESCRIPTION |  |

DESCRIPTION

This is a SOURCES SOUGHT NOTICE ONLY

This SOURCE SOUGHT NOTICE is for the purpose of Market Research.  Response to this notice will be used for information and planning purposes.  No proposals or quotes are being requested or accepted with this notice.  THIS IS NOT A SOLICITATION FOR PROPOSALS or QUOTES and NO CONTRACT SHALL BE AWARDED FROM THIS NOTICE.  This notice shall NOT be construed as solicitation or as an obligation on the part of the Department of Veterans Affairs (VA).  Responses to this notice is not a request to be added to a prospective bidders list or to receive a copy of the solicitation.

The Network Contracting Office (NCO) 16 is seeking preferably Service-Disabled Veteran Owned Small Businesses, Veterans Owned Small Businesses, or Small Businesses (but will accept other than small business for market research purposes), capable of providing Thyroid Uptake system and Neck Phantom Software to maintain and provide patient care on behalf of the Diagnostic Radiology Service for the Southeast Louisiana VHCS, New Orleans, Louisiana, per the DRAFT statement of work below.

The NAICS code is 334516 Analytical Laboratory Instrument manufacturing, and the Size Standard is 1000 employees.

Please provide a capability statement as well as any relevant literature addressing your organization’s ability to supply the required items. Please additionally provide a product specification sheet for the items.  Must provide an authorized distributor letter if not MFR.  Also, please include a point of contact name & phone number within email.

Responses must be received by no later than 10:00 a.m. Central Standard Time, April 27, 2026.  Responses shall be emailed to  tiffany.conner@va.gov.  Capabilities Statements to include UIE number.  Facsimile or Telephonic responses will not be accepted.

Note: All interested parties are reminded to be registered, active and verified in the following to be considered for Award.

SAM:   https://sam.gov/

If small business: SBA https://search.certifications.sba.gov/

(brand name or equivalent to)

Please confirm if you have a GSA/FSS schedule for these items.

Statement of Work

Thyroid Uptake System

Southeast Louisiana Veterans Health Care System

New Orleans, LA

1.	Purpose

1.1	The purpose is to purchase One (1) Thyroid Uptake system and One (1) Neck Phantom Software to maintain and provide patient care on behalf of the Diagnostic Radiology Service, for the Health Care System (SLVHCS) located at 2400 Canal St, New Orleans, LA 70119. The acquisition of the Captus 4000e will contribute significantly to the quality and efficiency of thyroid cancer treatment for veterans, ensuring they receive the best possible care with advanced medical technology.

2.0	Scope

2.1	The Contractor shall provide, transport, in-service and test all listed equipment.  All products must meet all salient characteristics defined in this section. All equipment shall be universal and approved for use by the FDA.

2.2	All equipment and installation must meet manufacturers, FDA and VA specifications.

2.3	The Contractor shall furnish all supplies, equipment, facilities, and services required for delivery and installation of the supplies and equipment.

2.4	The Contractor is responsible for any missing parts and components not included to carry out the installation.

2.5 Contractor must provide on-site service regarding but not limited to: installation and repairs.

2.6 All equipment must have a letter of validation.

2.8	Salient Characteristics

2.5.1 	THYROID UPTAKE SYSTEM

Quantity: 1 each

BRAND NAME or EQUIVALENT: Captus

Stock No. 5430-30154

Specifications/Salient Characteristics:

- Must have a 2 inch well

- Must include a color touch screen that provides a user interface

- Must include a Large 20-inch all-in-one touch screen computer

  - Touch Screen Flat Panel Computer

- Must be fully integrated DICOM interface for patient data

- Must include color Ink Jet printer

- Must have four heavy-duty locking casters that swivel 360°

- Must include Custom Protocol included as standard

- Must have an Ergonomically designed mobile stand

- Must include an  Articulating arm

  - 34-inch vertical range of travel

  - 220-degree collimator rotation

- Must include Medical-grade Corian countertop

- Must include a document storage shelf

- Must include Microsoft Office Professional 2010

- Must include a fully functional 1024 channel MCA with

auto and manual calibration.

- Must be able to create reports, including color graphics and isotope spectrum.

- Must Include: Isolation Transformer with hospital grade cord and plug

- Must have a protective cable management system

- Must have Fully protected internally routed cables

- Must be able to support Voltage

  - • 120 V ac,

  - 60 Hz, 2.6 AMP

  - 230 V ac

  - 50 Hz

  - 1.1 AMP

- Must include an uptake detector with the following specifications:

  - 2 in. Diameter 14-pin NaI(Tl) Detector:

    - Sodium Iodide detector that is thallium- doped

    - 14- pin connector

  - 1.55 in. Deep and 0.66 in. Diameter Well that measures the radioactive uptake during thyroid examinations

  - Flat Field Collimator

    - Must meet American National Standard Institutes

    - Must meet International Atomic energy agency requirements

  - Must be able to precisely measure and align the detector for precise distance from Thyroid

  - 1 inch and 2 inch well shield with brass liner

- Must include Detector inputs

- Channel Capacity

  - Must feature 1024 channels for spectral analysis

- Must be able to perform a maximum count rate of 200,000 counts per second at a rate of 5 percent

- Must be able to conduct count rate linearity with in 2 percent count rates of up to 150,000 cps for accurate data range intensities

- Must be able to set regions of interest that can be set either manually of automatically for specific areas on the spectrum

- Must be able to perform Differential Linearity that is less than 2 percent over the top of 98 percent that minimizes deviations during peak measurements.

- Must have the capability to preset for live time and conduct total counts allowing flexible peak measurements.

- Must include software controlled lower-level discriminator for setting lower threshold detection.

- Must weigh no more than 600 pounds

2.5.2	NECK PHANTOM FOR THYROID UPTAKE SYSTEM SOFTWARE

Quantity: 1 Each

BRAND NAME or EQUIVALENT: Captus

Stock no. 5230-0038

Specifications/Salient Characteristics:

- Must be able to accurately stimulate the human neck

- Must be able to ensure integration for optimal performance

- Must be able to provide consistent and reproducible positioning for calibrated procedures

4.	delivery

4.1	Contractor shall deliver all equipment to the Southeast Louisiana Veterans Health Care System (SLVHCS) Surgery Services 2400 Canal St, New Orleans, LA 70119

4.2	Deliver materials to job in manufacturer's original sealed containers with brand name marked thereon.

4.3	Package to prevent damage or deterioration during shipment, handling, and storage.

4.4	Maintain protective covering in place and in good repair until delivery is necessary.

4.5	Any government requested delayed delivery up to 90 days after initial award delivery date, shall be at no additional cost to the Government.

4.6	A pre-delivery meeting will be conducted 3 days prior to initial award delivery date for verification of delivery and installation dates.

4.7	SITE CONDITIONS

4.8	There shall be no smoking, eating, or drinking inside the hospital at any time.

5.	INSPECTION AND ACCEPTANCE:

5.1	The Contractor shall conduct a joint inspection with the COR upon delivery of tools.

5.2	Contractor shall provide dates of completion of punch list items and replacement parts and/or short ship items from the manufacturer(s).

6.	DELIVERABLES

6.1	Operation and Maintenance Manuals

6.1.1	Manuals - Quantity (2) each for items 2.5.1

6.2	Deliver compilation of all manufacturers recommended maintenance schedule and operation materials to COR upon completion of installation.

7.	TRAINING:

7.1	Contractor shall provide On-site training of the equipment to the Users.  Scheduling of operator training shall be coordinated with the SLVHCS COR after installation is complete.

7.2	The contractor shall provide a training program to the Biomedical Engineering technicians and in-house clinical personnel.  Technical training must provide Biomedical Engineering with the tools and knowledge to fully operate and maintain the system.