## DLSS Service Parts List
EODGRU1 DIVERS LIFE SUPPORT SYSTEMS (DLSS) MAINTENANCE SERVICES SERVICE PARTS LIST
ITEM NO. | ITEM DESCRIPTION | QTY | COMMENTS | COMMENTS 2
1 | Gas Cylinder Fill x 100 | 50 ea | Refill HeO2 88/12% compressed gas cylinders (certificate of gas analysis included)
2 | Gas Cylinder Fill x 100 | 30 ea | Refill O2 100% compressed gas cylinders (certificate of gas analysis included)
3 | Gas Cylinder Fill x 100 | 20 ea | Refill O2N2 60/40% compressed gas cylinders (certificate of gas analysis included)
4 | Gas Cylinder Fill x 100 | 10 ea | Various cylinder fill fees (pick up charge, delivery charge, fuel surcharge, hazardous material charge, driver retention charge)
5 | OTPA Maintenance x 3 | 3 ea | 24M Calibrate oxygen outlet pressure gauge OHP-G201 | 24M indicates maintenance service is required every 24 months.
6 | OTPA Maintenance x 3 | 3 ea | 24M Replace Pump Assembly Air Cycling Valve O-rings and Bumper | 24M indicates maintenance service is required every 24 months.
7 | OTPA Maintenance x 3 | 3 ea | Provide air cycling valve kit (p/n 17329)
8 | OTPA Maintenance x 3 | 3 ea | 36M Test Oxygen Transfer Pump Assembly Relief Valves | 36M indicates maintenance service is required every 36 months.
9 | OTPA Maintenance x 3 | 3 ea | 36M Inspect Oxygen In-Line Filters | 36M indicates maintenance service is required every 36 months.
10 | OTPA Maintenance x 3 | 3 ea | 72M Deliver Oxygen Transfer Pump Assembly for Overhaul | 72M indicates maintenance service is required every 72 months.
11 | Bauer O26-D Maintenance x 4 | 2 ea | A-2 Replace Air Purifier (SECURUS) Cartridge, Drying Filter Cartridge, and CO Filter Cartridge
12 | Bauer O26-D Maintenance x 4 | 2 ea | Provide drying filter cartridge (p/n 058825A)
13 | Bauer O26-D Maintenance x 4 | 8 ea | Provide Securus filter cartridge (p/n 060037A)
14 | Bauer O26-D Maintenance x 4 | 16 ea | Provide O-ring (p/n N04735)
15 | Bauer O26-D Maintenance x 4 | 16 ea | Provide backup ring (p/n N04736)
16 | Bauer O26-D Maintenance x 4 | 4 ea | Provide air intake filter (p/n N25886)
17 | Bauer O26-D Maintenance x 4 | 4 ea | 18M Calibrate Diver's Portable Gauges | 18M indicates maintenance service is required every 18 months.
18 | Bauer O26-D Maintenance x 4 | 20 hrs | 18M Disassemble, Clean, Inspect and Reassemble Sintered Metal Filter or Helical Insert Bauer Oil and Water Separator Element
19 | Bauer O26-D Maintenance x 4 | 4 ea | Provide sintered filter element (p/n 061860) | 18M indicates maintenance service is required every 18 months.
20 | Bauer O26-D Maintenance x 4 | 2 ea | 24M Change Oil and Replace Oil Tube/Filter Assembly | 24M indicates maintenance service is required every 24 months.
21 | Bauer O26-D Maintenance x 4 | 2 ea | Provide oil filter (p/n N25326)
22 | Bauer O26-D Maintenance x 4 | 2 ea | Provide gasket (p/n 077774)
23 | Bauer O26-D Maintenance x 4 | 2 ea | Provide O-ring (p/n N04058)
24 | Bauer O26-D Maintenance x 4 | 2 ea | Provide O-ring (p/n N25327)
25 | Bauer O26-D Maintenance x 4 | 8 gal | Provide compressor oil (p/n OIL-0024)
26 | Bauer O26-D Maintenance x 4 | 4 ea | 24M Replace Bauer Air Compressor Pressure Maintaining Valve (PMV) and Inlet/Outlet Check Valves | 24M indicates maintenance service is required every 24 months.
27 | Bauer O26-D Maintenance x 4 | 4 ea | Provide check valve (p/n VAL-0007)
28 | Bauer O26-D Maintenance x 4 | 8 ea | Provide pressure maintaining valve (p/n VAL-0053)
29 | Bauer O26-D Maintenance x 4 | 4 ea | 36M Test Relief Valves Removed From System | 36M indicates maintenance service is required every 36 months.
30 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | A-10 Replace Bauer air purifier element
31 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | Provide air purifier element (p/n 058821A)
32 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | Provide O-ring (p/n N04735)
33 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | Provide backup ring (p/n N04736)
34 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | A-11 Change CO/Moisture indicator element
35 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | Provide moisture indicator element (p/n ELM-0056)
36 | C-D/DV/Navy Compressor Maintenance x 5 | 10 ea | 18M Calibrate Diver's Portable Gauges | 18M indicates maintenance service is required every 18 months.
37 | C-D/DV/Navy Compressor Maintenance x 5 | 5 ea | 36M Test Relief Valves Removed From System | 36M indicates maintenance service is required every 36 months.
38 | Transportable Recompression Chamber System (TRCS) x 3 | 72 ea | 24M Calibrate Diver's Life Support System (DLSS) Installed Gauges onsite | 24M indicates maintenance service is required every 24 months.
39 | Transportable Recompression Chamber System (TRCS) x 3 | 18 ea | Labor involved in TRCS calibration (relief valve removal, system line-ups, joint testing)
40 | Transportable Recompression Chamber System (TRCS) x 3 | 18 ea | Provide AS3581-111 O-rings, cleaned IAW MIL STD 1330D (for TRCS relief valve removal)
41 | Transportable Recompression Chamber System (TRCS) x 3 | 5 ea | Provide AS3581-010 O-rings, cleaned IAW MIL STD 1330D (for TRCS med-lock gauge)
42 | Transportable Recompression Chamber System (TRCS) x 3 | 3 ea | 36M Perform Visual Inspection of Acrylic Window | 36M indicates maintenance service is required every 36 months.
43 | Transportable Recompression Chamber System (TRCS) x 3 | 24 ea | 36M Test Relief Valves Removed from System onsite | 36M indicates maintenance service is required every 36 months.
44 | Transportable Recompression Chamber System (TRCS) x 3 | 18 hrs | Labor involved in TRCS calibration (relief valve removal, system line-ups, joint testing)
45 | Transportable Recompression Chamber System (TRCS) x 3 | 18 ea | Provide AS3581-111 O-rings, cleaned IAW MIL STD 1330D (for TRCS relief valve removal)
46 | Transportable Recompression Chamber System (TRCS) x 3 | 15 hrs | 36M Clean and Inspect Transportable Recompression Chamber (TRC) T-Filter Elements and Housings | 36M indicates maintenance service is required every 36 months.
47 | Transportable Recompression Chamber System (TRCS) x 3 | 15 ea | Provide Swagelok filter element (p/n M-8F-K4-15)
48 | Transportable Recompression Chamber System (TRCS) x 3 | 15 ea | Provide Swagelok filter gasket (p/n SS-8TF-K2)
49 | Transportable Recompression Chamber System (TRCS) x 3 | 6 hrs | Clean filter elements and gaskets IAW MIL STD 1330E
50 | Transportable Recompression Chamber System (TRCS) x 3 | 12 hrs | 36M Clean and Inspect Transfer Lock (TL) T-Filter Elements and Housings | 36M indicates maintenance service is required every 36 months.
51 | Transportable Recompression Chamber System (TRCS) x 3 | 12 ea | Provide Swagelok filter element (p/n M-8F-K4-15)
52 | Transportable Recompression Chamber System (TRCS) x 3 | 12 ea | Provide Swagelok filter gasket (p/n SS-8TF-K2)
53 | Transportable Recompression Chamber System (TRCS) x 3 | 6 hrs | Clean filter elements and gaskets IAW MIL STD 1330E
54 | Transportable Recompression Chamber System (TRCS) x 3 | 36 ea | 60M Perform DOT .93 Carbon Fiber Flask(s) External and Internal Inspection, Hydrostatic Testing and Rupture Disc Replacement | 60M indicates maintenance service is required every 60 months.
55 | Transportable Recompression Chamber System (TRCS) x 3 | 36 hrs | Labor to disassemble and reassemble (4) flask racks assemblies while maintaining cleanliness IAW MIL STD 1330E, including filling flasks for 24 hour joint tightness drop test
56 | Transportable Recompression Chamber System (TRCS) x 3 | 36 ea | Provide rupture discs (p/n 20Z), cleaned IAW MIL STD1330E
57 | Transportable Recompression Chamber System (TRCS) x 3 | 36 ea | Provide M83248/2-916 o-rings, cleaned IAW MIL STD 1330E
58 | Transportable Recompression Chamber System (TRCS) x 3 | 36 ea | Provide M83248/2-010 O-rings, cleaned IAW MIL STD 1330E
59 | Transportable Recompression Chamber System (TRCS) x 3 | 3 ea | Diver’s air gas required for flask rack joint tightness test (leave flasks full and provide current satisfactory air sample from diver’s air source)
60 | Transportable Recompression Chamber System (TRCS) x 3 | 3 ea | 60M Hydrostatically Test Fire Extinguisher | 60M indicates maintenance service is required every 60 months.
61 | Transportable Recompression Chamber System (TRCS) x 3 | 15 hrs | U-2 Hydrostatically Test Hoses
62 | Transportable Recompression Chamber System (TRCS) x 3 | 24 hrs | U-3 Clean and Inspect Interlock Poppet and NATO Flange Bearings
63 | Transportable Recompression Chamber System (TRCS) x 3 | 9 hrs | U-5 Clean, Inspect and Reinstall or Replace Air Supply and Exhaust Diffusers/Silencers
64 | Transportable Recompression Chamber System (TRCS) x 3 | 12 hrs | REC and OQE package
65 | SNDL Maintenance x 1 | 1 ea | 18M Test and Inspect ASRA Check Valve (AHP-V445). | 18M indicates maintenance service is required every 18 months.
66 | SNDL Maintenance x 1 | 22 ea | 24M Calibrate Diver's Life Support System (DLSS) Installed Gauges. | 24M indicates maintenance service is required every 24 months.
67 | SNDL Maintenance x 1 | 2 hrs | 36M Inspect Standard Navy Double Lock Recompression Chamber System (SNDLRCS) HP In-Line Air Reducing Station Filter Element | 36M indicates maintenance service is required every 36 months.
68 | SNDL Maintenance x 1 | 2 ea | Provide filter element (p/n U-10007)
69 | SNDL Maintenance x 1 | 2 hrs | Clean (2) filter elements IAW MIL STD 1622B
70 | SNDL Maintenance x 1 | 2 ea | Provide M83248/2-906 O-rings, cleaned IAW MIL STD 1622B
71 | SNDL Maintenance x 1 | 2 ea | Provide M83248/2-914 O-rings, cleaned IAW MIL STD 1622B
72 | SNDL Maintenance x 1 | 2 ea | Provide M83248/2-015 O-rings, cleaned IAW MIL STD 1622B
73 | SNDL Maintenance x 1 | 4 ea | Provide M83248/2-908 O-rings, cleaned IAW MIL STD 1622B
74 | SNDL Maintenance x 1 | 4 hrs | 36M Clean, Inspect and Test Standard Navy Double Lock (SNDL) HP Oxygen Reducing Station Piping and Valves. | 36M indicates maintenance service is required every 36 months.
75 | SNDL Maintenance x 1 | 10 hrs | Labor involved in SNDL HP oxygen reducing station valve inspection (panel removal, piping removal/reinstallation, panel reinstallation, joint testing)
76 | SNDL Maintenance x 1 | 1 hr | Visually inspect primary HP oxygen supply regulator OHP-R206 and secondary HP oxygen supply regulator OHP-R205 inlet screens
77 | SNDL Maintenance x 1 | 1 hr | Visually inspect primary oxygen supply valve OHP-V201 and secondary oxygen supply valve OHP-V203 outlets
78 | SNDL Maintenance x 1 | 4 hrs | Overhaul primary oxygen supply valve OHP-V201 and secondary oxygen supply valve OHP-V203
79 | SNDL Maintenance x 1 | 2 ea | Provide valve repair kit (p/n T-9K-6N)
80 | SNDL Maintenance x 1 | 2 hr | Clean (2) valve repair kits IAW MIL STD 1330E
81 | SNDL Maintenance x 1 | 1 hr | Perform valve seat tightness test on primary oxygen supply valve OHP-V201 and secondary oxygen supply valve OHP-V203
82 | SNDL Maintenance x 1 | 7 ea | 36M Test Relief Valves Removed from System | 36M indicates maintenance service is required every 36 months.
83 | SNDL Maintenance x 1 | 3 ea | 36M Clean and Inspect Standard Navy Double Lock Recompression Chamber System (SNDLRCS) T-Filter Elements and Housings | 36M indicates maintenance service is required every 36 months.
84 | SNDL Maintenance x 1 | 3 ea | Provide Swagelok filter element (p/n M-8F-K4-15)
85 | SNDL Maintenance x 1 | 3 ea | Provide Swagelok filter gaskets (p/n SS-8TF-K2)
86 | SNDL Maintenance x 1 | 3 hr | Clean (3) filters and (3) gaskets IAW MIL STD 1330E
87 | SNDL Maintenance x 1 | 3 hrs | 36M Clean and Inspect Standard Navy Double Lock Recompression Chamber System (SNDLRCS) LP In-Line Filter Element | 36M indicates maintenance service is required every 36 months.
88 | SNDL Maintenance x 1 | 2 ea | Provide Norman filter element (p/n 416G-20VN)
89 | SNDL Maintenance x 1 | 2 hrs | Clean (2) filter elements IAW MIL STD 1622B
90 | SNDL Maintenance x 1 | 4 ea | Provide M83248/2-924 O-rings, cleaned IAW MIL STD 1622B
91 | SNDL Maintenance x 1 | 4 ea | Provide M83248/2-912 O-rings, cleaned IAW MIL STD 1622B
92 | SNDL Maintenance x 1 | 8 ea | Provide M83248/2-021 O-rings, cleaned IAW MIL STD 1622B
93 | SNDL Maintenance x 1 | 2 hrs | 36M Remove, Clean, Inspect, Re-Use or Replace Charge Port and HP Regulator Filters of ASRA/OSRA/HOSRA and SCUBA | 36M indicates maintenance service is required every 36 months.
94 | SNDL Maintenance x 1 | 2 ea | Provide filter element (p/n U-10007)
95 | SNDL Maintenance x 1 | 2 hrs | Clean (2) filter elements IAW MIL STD 1622B
96 | SNDL Maintenance x 1 | 4 ea | Provide M83248/2-012 O-rings, cleaned IAW MIL STD 1622B
97 | SNDL Maintenance x 1 | 4 ea | Provide M83248/2-906 O-rings, cleaned IAW MIL STD 1622B
98 | SNDL Maintenance x 1 | 2 ea | Provide M83248/2-914 O-rings, cleaned IAW MIL STD 1622B
99 | SNDL Maintenance x 1 | 9 ea | 60M Perform DOT 3.15 Carbon Fiber Flasks External and Internal Inspection, Hydrostatic Testing and Rupture Disc Replacement | 60M indicates maintenance service is required every 60 months.
100 | SNDL Maintenance x 1 | 16 hrs | Labor to disassemble, remove, reassemble, and reinstall flasks while maintaining cleanliness, including filling flasks for 24 hour joint tightness drop test
101 | SNDL Maintenance x 1 | 9 ea | Provide rupture discs (p/n 128450XX)
102 | SNDL Maintenance x 1 | 9 hrs | Clean (9) rupture discs IAW MIL STD 1622B
103 | SNDL Maintenance x 1 | 15 ea | Provide M83248/2-111 o-rings, cleaned IAW MIL STD 1330D
104 | SNDL Maintenance x 1 | 18 ea | Provide M83248/2-910 o-rings, cleaned IAW MIL STD 1330D
105 | SNDL Maintenance x 1 | 17 ea | Provide M83248/2-010 o-rings, cleaned IAW MIL STD 1330D
106 | SNDL Maintenance x 1 | 1 ea | Diver’s air gas required for joint tightness/24hr drop test (leave flasks full and provide current satisfactory air sample from diver’s air source)
107 | SNDL Maintenance x 1 | 10 hrs | U-2 Hydrostatically Test Hose
108 | SNDL Maintenance x 1 | 2 hrs | U-4 Repair or Replace ASRA Check Valve (AHP-V445) and Conduct Tightness Tests
109 | SNDL Maintenance x 1 | 1 ea | Provide Swagelok check valve (p/n SS-CHVCO8-5)
110 | SNDL Maintenance x 1 | 1 hr | U-5 Clean, Inspect and Reinstall or Replace Air Supply and Exhaust Diffuser(s)/Silencer
111 | SNDL Maintenance x 1 | 4 hrs | REC and OQE package
112 | FADS III Maintenance x 1 | 12 ea | 24M Calibrate Diver's Life Support System (DLSS) Installed Gauges onsite | 24M indicates maintenance service is required every 24 months.
113 | FADS III Maintenance x 1 | 1 hrs | 36M Clean and Inspect Control Console Assembly (CCA) HP In-Line Air Filter Elements | 36M indicates maintenance service is required every 36 months.
114 | FADS III Maintenance x 1 | 1 ea | Provide filter element (p/n U-10007)
115 | FADS III Maintenance x 1 | 4 hrs | Clean (4) filter elements IAW MIL STD 1622B
116 | FADS III Maintenance x 1 | 6 ea | Provide M83248/2-906 O-rings, cleaned IAW MIL STD 1622B
117 | FADS III Maintenance x 1 | 1 ea | Provide M83248/2-914 O-rings, cleaned IAW MIL STD 1622B
118 | FADS III Maintenance x 1 | 2 ea | Provide M83248/2-111 O-rings, cleaned IAW MIL STD 1622B
119 | FADS III Maintenance x 1 | 4 ea | 36M Test Relief Valves Removed from System onsite | 36M indicates maintenance service is required every 36 months.
120 | FADS III Maintenance x 1 | 1 hrs | 36M Remove, Clean, Inspect, Re-Use or Replace Charge Port and HP Regulator Filters of ASRA/OSRA/HOSRA and SCUBA | 36M indicates maintenance service is required every 36 months.
121 | FADS III Maintenance x 1 | 1 ea | Provide filter element (p/n U-10007)
122 | FADS III Maintenance x 1 | 4 hrs | Clean (4) filter elements IAW MIL STD 1622B
123 | FADS III Maintenance x 1 | 2 ea | Provide M83248/2-012 O-rings, cleaned IAW MIL STD 1622B
124 | FADS III Maintenance x 1 | 1 ea | Provide M83248/2-906 O-rings, cleaned IAW MIL STD 1622B
125 | FADS III Maintenance x 1 | 1 ea | Provide M83248/2-914 O-rings, cleaned IAW MIL STD 1622B
126 | FADS III Maintenance x 1 | 9 ea | 60M Perform DOT 3.15 Carbon Fiber Flasks External and Internal Inspection, Hydrostatic Testing and Rupture Disc Replacement | 60M indicates maintenance service is required every 60 months.
127 | FADS III Maintenance x 1 | 16 hrs | Labor to disassemble, remove, reassemble, and reinstall flasks while maintaining cleanliness, including filling flasks for 24 hour joint tightness drop test
128 | FADS III Maintenance x 1 | 9 ea | Provide rupture discs (p/n 128450XX)
129 | FADS III Maintenance x 1 | 9 hrs | Clean (9) rupture discs IAW MIL STD 1622B
130 | FADS III Maintenance x 1 | 15 ea | Provide M83248/2-111 o-rings, cleaned IAW MIL STD 1330D
131 | FADS III Maintenance x 1 | 18 ea | Provide M83248/2-910 o-rings, cleaned IAW MIL STD 1330D
132 | FADS III Maintenance x 1 | 17 ea | Provide M83248/2-010 o-rings, cleaned IAW MIL STD 1330D
133 | FADS III Maintenance x 1 | 1 ea | Diver’s air gas required for joint tightness/24hr drop test (leave flasks full and provide current satisfactory air sample from diver’s air source)
134 | FADS III Maintenance x 1 | 5 hrs | U-2 Hydrostatically Test Hoses
135 | FADS III Maintenance x 1 | 4 hrs | REC and OQE package
136 | ORCA Maintenance x 1 | 2 ea | 24M Calibrate Diver's Life Support System (DLSS) Installed Gauges | 24M indicates maintenance service is required every 24 months.
137 | ORCA Maintenance x 1 | 1 ea | 36M Test Relief Valves Removed From System | 36M indicates maintenance service is required every 36 months.
138 | ORCA Maintenance x 1 | 3 hrs | 36M Clean and Inspect Oxygen Regulating Console Assembly (ORCA) I & II Inline Filter Elements | 36M indicates maintenance service is required every 36 months.
139 | ORCA Maintenance x 1 | 2 hrs | U-2 Hydrostatically Test Hoses
140 | ORCA Maintenance x 1 | 2 hrs | REC and OQE package
141 | Relief Valves x 200 | 200 ea | 36M Test Relief Valves Removed from System onsite | 36M indicates maintenance service is required every 36 months.
142 | SCUBA Cylinders x 300 | 300 ea | 60M Hydrostatic test SCUBA cylinders | 60M indicates maintenance service is required every 60 months.
143 | DLSS Gauges x 400 | 200 ea | 18M 6R Calibrate Diver's Life Support System (DLSS) Portable Gauges | 18M indicates maintenance service is required every 18 months.
144 | DLSS Gauges x 400 | 200 ea | 24M Calibrate Diver's Life Support System (DLSS) Installed Gauges | 24M indicates maintenance service is required every 24 months.
145 | Closed Circuit UBA Viper x 2 | 2 ea | 24M Calibrate Test Panel Gages, Relief Valve, and Conduct Comparison Accuracy Test for Flowmeters and Pocket Flowmeters. | 24M indicates maintenance service is required every 24 months.
146 | Closed Circuit UBA Viper x 2 | 1 hrs | 60M Deliver Pressure Vessels to an Authorized Activity for Hydrostatic Testing. | 60M indicates maintenance service is required every 60 months.
Note on B3: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 100 indicates there are 100 gas cylinder fills.
Note on C3: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 50 each indicates each gas cylinder requires 50 gas refills.
Note on B13: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 4 indicates there are 4 Bauer O26-Ds.
Note on C13: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 2 each indicates each Bauer compressor requires replacement 2 times.
Note on C27: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: 8 gal indicates that 8 gallons of oil are required for each Bauer compressor.